3701:1-66-04     Quality assurance program for medical radiation-generating equipment.

(A) Each registrant will develop, implement and maintain a written quality assurance program in the form of a readily available manual or manuals, either in hard copy format or electronic format. For purposes of this chapter and Chapter 3701:1-67 of the Administrative Code, quality assurance program means a program providing for verification by written procedures such as testing, auditing, and inspection to ensure that deficiencies, deviations, defective equipment, or unsafe practices, or a combination thereof, relating to the use, disposal, management, or manufacture of radiation devices are identified, promptly corrected, and reported to the appropriate regulatory authorities.

Each registrant must develop a written quality assurance program (policies and procedures) in the form of readily available manual or manuals, either in hard copy or electronic format, that identifies sound radiation safety practices and promotes the safe operation of radiation-generating equipment.

A well written quality assurance program promotes prompt identification, correction and reporting (if necessary) of deficiencies, deviations, defective equipment or unsafe practices.

This written quality assurance program must be implemented - meaning that the quality assurance program is being followed by the staff.

This written quality assurance program must be maintained – meaning the quality assurance program must be revised, as necessary, to reflect any changes to radiation safety practices and organized in a written or electronic document.

(B) The written quality assurance program of each registrant will address and include records to verify implementation of at least the following:

The written quality assurance program must contain policies and procedures as described in this rule.

Records to verify implementation shall include documenting results of the tests, audits and inspections performed by responsible facility personnel.

(1) The intervals of and procedures for the evaluation of all radiation-generating equipment to ensure compliance with all applicable rules of this chapter;

The written quality assurance program must indicate a specific interval at which the radiation-generating equipment is evaluated and calibrated to ensure that the radiation-generating equipment is working according to all applicable requirements of Chapter 3701:1-66 of the Ohio Administrative Code.

The procedures for such an evaluation may include periodic evaluation by a radiation expert (medical physicist), the performance of manufacturer recommended preventative maintenance by appropriately trained personnel, or a combination of the above.

Typical compliance evaluations include testing the reproducibility and accuracy of the radiation output, the accuracy of the beam restriction device (collimators) and the condition of the x-ray beam, according to clinical use factors.

NOTE: The periodic inspection by the Ohio Department of Health does not meet the evaluation and calibration requirements of this rule. The Ohio Department of Health’s inspection is only to verify compliance with the rule requirements.

(2) Procedures for maintaining compliance with occupational and public exposure limits;

The written quality assurance program must include procedures to address compliance with occupational and public exposure limits in rules 3701:1-38-12 and 3701:1-38-13 of the Ohio Administrative Code (OAC).

Refer to rules 3701:1-38-14(B) and 3701:1-66-02(G)(5) of the OAC to determine if personnel monitoring (badges) are necessary.

- If unnecessary, the written quality assurance program will identify that personnel monitoring is not required and all individuals are to follow the time, distance, and shielding requirements of the quality assurance program.

- If personnel monitoring is necessary or provided, the quality assurance program will identify the following:

  - Rule 3701:1-38-14(C) of the OAC, wearing individual monitoring devices (badges): For example:

- Whole body monitoring, the individual monitoring device must be worn at an unshielded location likely to receive the highest exposure. When an apron is worn, the location is usually the neck.

- Individual monitoring devices for monitoring dose to the embryo or fetus of a declared pregnant women pursuant to paragraph (H) of rule 3701:1-38-12 of the OAC shall be located at the waist under any protective apron being worn the woman.

  - Rule 3701:1-38-10(C) of the OAC, requirements to provide monitoring records to individuals:

- Annually – if the individual’s occupational dose exceeds one millisievert (one hundred millirem) TEDE or one millisievert (one hundred millirem) to any individual organ or tissue.

- At the request of a worker formerly engaged in radiation activities controlled by the facility.

- When a facility is required pursuant to paragraphs (A) to (C) of rule 3701:1-38-21 of the OAC to report any exposure of an individual to the Director, the facility shall also provide the individual a written report on the exposure data included in the report to the Director.

- At the request of an employee who is terminating employment.

(3) Procedures for notifying the director when individuals are occupationally over-exposed to radiation, pursuant to Chapter 3701:1-38 of the Administrative Code;

If individual monitoring devices (badges) are required or provided, the written quality assurance program of the facility must include a procedure for notifying the Director of individuals occupationally over-exposed to radiation.

The procedure must identify that the Director will be notified in accordance with rule 3701:1-38-21 of the OAC.

(4) Safe operating procedures for each type of radiation-generating equipment to be handled;

The written quality assurance program must address safe operating procedures. The safe operating procedures must address each type of radiation-generating equipment on site. If the facility has several radiographic units that are of the same type, then only one set of safe operating procedures for those units is necessary.

The safe operating procedures may include topics such as but not limited to:

- Follow manufacture specification for equipment operation, morning quality control if applicable and warm-up

- Use devices or administrative procedures to prevent unauthorized use of the radiation-generating equipment 3701:1-38-17

- Position of operator during radiation exposures

- Six feet from dental I/O unit when switch is on a cord

- Behind protective barrier and able to see the patient when exposure switch is permanently mounted behind a protective barrier

- Use appropriate techniques for patient examinations from a technique chart or preprogramed control panel 3701:1-66-02(G)(1)

- Proper use of the radiation-generating equipment’s collimation to reduce radiation exposure

- Proper use of mobile barriers, lead aprons and lead gloves

- Provide appropriate protective apparel for individuals required for the radiographic procedure 3701:1-66-02(G)(2)

- Only appropriate licensed individuals may operate radiation-generating equipment to expose human beings. 3701:1-66-02(C)(1)

- When multiple tubes are controlled by one exposure switch, the operator must know how to tell which tube is selected 3701:1-66- 02(F)(5)

- If applicable, proper use of a manual line voltage compensator to maintain appropriate energy

- Indication that only licensed practitioners order x-ray examinations 3701:1-66-02(B) and 3701:1-66-04(B)(17).

- Proper attachment of removable spacer cone on c-arm fluoroscopic equipment 3701:1-66-07(A)(2)

- If applicable, limitations on source-to-image distance selection due to grid focal length

- The radiation-generating equipment tolerance limits and not to exceed those limits

(5) Training of operators of each type of radiation-generating equipment to be handled in order to assure competency in the operating procedures;

The written quality assurance program must address operator training specific to each type of x-ray unit at the facility.

The records for verifying implementation of the operator training program must be provided for anyone who takes x-rays at the facility.

The inspector will review the quality assurance policy and the training records of all individuals that have operated radiation-generating equipment during the inspection period.

(6) In addition to the obligations of paragraph (B)(1) of rule 3701:1-38-10 of the Administrative Code, individuals likely to receive an annual occupational dose in excess of one millisievert (one hundred millirem) will be instructed in the following:

This rule is applicable to individuals that are likely to receive an annual occupational dose in excess of one millisievert (one hundred millirem).

Individuals that work in a restricted area should receive instruction unless it has been historically documented that the individuals will not receive an annual occupational dose in excess of 100 mR or a Radiation Expert has determined the dose will not be exceeded. Generally, any facility that is required to badge their employees should provide this instruction.

Instruction must be documented.  Typically, signatures are acceptable evidence that instruction occurred.

Dental facilities do not meet the occupational dose limits discussed in this rule and therefore are not required to provide this instruction.

(a) The location, boundaries, and purpose of restricted areas; and

The written quality assurance program must describe the restricted area(s):

- A restricted area is an area limited by the facility for the purpose of protecting individuals against undue risks from exposure to radiation.

- Clear identification and enforcement of a location (such as a room) as being a restricted area during radiation exposure is acceptable.

- Clear identification of the boundaries of open areas as restricted areas must also have methods for operator enforcement to assure those boundaries are maintained during radiation exposure.

NOTE: the universal three-blade radiation signs are good for clearly outlining boundaries or restricted areas.

(b) A description of the radiation-generating equipment and its location;

The written quality assurance program must describe the type of radiation-generating equipment and its location:

Examples:

- An intraoral unit is in room A.

- Room A has an intraoral unit.

- Room A has a computed tomography unit.

(7) The quality control tests to be performed, the frequency of the quality control tests to be performed and the personnel responsible for the performance of the quality control tests as applicable to the radiation-generating equipment type and use;

The written quality assurance program must address quality control testing.

At a minimum, quality control testing procedures must address (as applicable):

- The temperature of their film developer chemistry [ref. rule 3701:1-66-02(J)(3)(a) or (b) of the Ohio Administrative Code (OAC)].

- Film processing solution preparation and maintenance [ref. rule 3701:1-66-02(J)(3)(d), OAC].

- Cassette/screen or cassette/computed radiography (phosphorus) imaging plate maintenance [ref. rule 3701:1-66-02(J)(9) of the OAC].

- The process that will ensure computed radiography (phosphorus) plates are used frequently enough to produce diagnostic quality images [ref. rule 3701:1-66-02(J)(10)(c) of the OAC and manufacturer specifications].

- [Fluoroscopy equipment only] Initial and annual quality control tests that must be performed by a radiation expert (medical physicist) [ref. 3701:1-66-07(C)(2)(a), (C)(2)(b) and (C)(2)(c) of the OAC].

- [Stereotactically-guided breast biopsy equipment only] Initial and annual quality control tests that must be performed by a radiation expert (medical physicist) [ref. 3701:1-66-08(G) of the OAC].

- [Computerized tomography (CT) equipment only]:

- [Conventional (fan beam) CT]:  Initial and annual quality control tests that must be performed by a radiation expert (medical physicist) [ref. 3701:1-66-10(C)(1)(a) (C)(1)(b) of the OAC]. Quality control program that must be performed by the CT technologist and must be approved by a radiation expert (medical physicist) [ref. 3701:1-66-10(C)(1) of the OAC].

- [Cone beam CT]: Quality control test procedures, frequencies and tolerance limits developed under the guidance of a radiation expert (medical physicist) [ref. 3701:1-66-10(D)(1), of the OAC]. Initial and annual quality control tests that must be performed by a radiation expert (medical physicist) [ref. 3701:1-66- 10(C)(1), of the OAC].

- [Bone densitometry equipment only] Quality control tests as specified by the manufacturer [ref. 3701:1-66-11(C)(1), OAC].

The identified responsible personnel may be from within the facility staff, from an outside company or a combination of the two. At a minimum, unless specified by another OAC rule, facility records must list responsible personnel by job description (a list by name may be appropriate but is not required to meet compliance with this rule).

(8) Policies regarding the state licensure or certification of each person operating radiation-generating equipment as obligated by Chapters 4773. and 4715. of the Revised Code;

The written quality assurance program must contain policies that address the process that will be used to ensure everyone who operates radiation-generating equipment holds a valid license or certificate as required by the Ohio Revised Code.

Example: Only individuals appropriately licensed in accordance with Chapter 3701-72 of the Ohio Administrative Code are allowed to take x-rays. The facility verifies active licenses via the Ohio Department of Health website.

(9) The dissemination of quality assurance policies and a method to educate affected workers on those policies and any policy changes;

The written quality assurance program must indicate:

- How quality assurance policies are provided to all workers that would be affected by those policies.

- The method used to provide policy changes to affected workers.

(10) Radiation workers' role and responsibility for following and supporting the quality assurance program;

The written quality assurance program must identify that the radiation worker is responsible for following and supporting the policies and procedures of the radiation quality assurance program. This support should include enforcing the radiation safety policies and procedures of the quality assurance program.

(11) Policies regarding personnel protection, including time, distance, and shielding;

The written quality assurance program must contain policies regarding time, distance and shielding for the purpose of personnel radiation protection.

Examples:

- Other than the patient no one is allowed in the x-ray room during radiation exposures.

- When participating in fluoroscopy procedures wear lead aprons and maintain as much distance as possible from the fluoroscopy tube and image intensifier.

- Personnel assisting during an x-ray procedure must be positioned such that no part of the body is struck by the useful beam and shall be protected from direct scatter radiation by protective aprons or whole-body protective barriers of not less than .25 millimeters of lead equivalent material.

(12) Policies regarding occupational exposure of pregnant workers;

The written quality assurance program must contain policies regarding the occupational exposure of pregnant workers.

At a minimum, the policy must specify that the facility will follow the requirements of rule 3701:1-38-10, 3701:1-38-12, 3701:1-38-14 and 3701:1-38-20 of the Ohio Administrative Code (OAC).

(13) Policies regarding radiation safety training for ancillary personnel;

‘Ancillary personnel’ are those individuals who are present in the room during the operation of radiation-generating equipment, but not directly involved in the operation. Such individuals may include nurses, surgeons, anesthesiology staff, surgical assistants, speech pathology staff, emergency department staff, etc. Under conditions where ancillary personnel are occupationally exposed to radiation, basic radiation safety training must be provided.

The written quality assurance program must include policies regarding the training of ancillary personnel.

(14) Policies regarding training for personnel with quality control responsibilities;

The written quality assurance program must include policies addressing training for personnel with quality control responsibilities. The policies may be specific to an individual or generally applied to all personnel with quality control responsibilities in the job description.

(15) Policies regarding human patient protection, including screening for pregnancy, exposure of pregnant patients, patient shielding, patient education;

The written quality assurance program must contain policies regarding the topics listed in this rule.

Policy example:

- Patient protection – Radiation signs posted to identify restricted areas. Sponges and positioning devices used to support patients during examinations.

- Pregnancy screening – Patient pregnancy will be verified verbally prior to the x-ray procedure. If a patient is unsure if they are pregnant a pregnancy test will be administered.

- Exposure of pregnant patient – Prior to the x-ray, the physician or the patient’s OBGYN will determine the benefit versus the risk of having the x-ray.

- Patient shielding – The facility policy to shield patients with lead aprons. NOTE that even though lead aprons are not required by rule for shielding patients’ gonadal region anymore, a policy on shielding is still required. The policy may be that shielding is no longer required, but it offered as an option to patients to make them feel more comfortable.

- Patient education – Written information provided to patients such as pamphlets describing the examination, preparation and post instructions, etc. Or a policy that a verbal explanation is provided.

(16) Policies regarding verification of human patient identity and exam to be performed, including identification of the appropriate body part;

The written quality assurance program must include policies addressing the verification of patient identity and exam to be performed.

Patient identification policy example – The individual performing the x-ray exam verifies patient identity (the best practice, is two methods, such as name and date of birth). The individual performing the x-ray exam verifies the exam and appropriate body part (right/left etc.) and ask the patient what x-ray exam they are having.

(17) Policies to only permit licensed practitioners to order radiographic examinations;

The written quality assurance program must include policies addressing the ordering of radiographic examinations.

A licensed practitioner is one of the following:

A Dentist licensed in Ohio pursuant to Chapter 4715. of the Revised Code;

A Physician licensed in Ohio to practice medicine or surgery or osteopathic medicine or surgery pursuant to Chapter 4731. of the Revised Code;

A Podiatrist licensed in Ohio pursuant to Chapter 4731. of the Revised Code;

A Veterinarian licensed in Ohio pursuant to Chapter 4741. of the Revised Code, can order x-ray on animals only;

A Chiropractor licensed in Ohio pursuant to Chapter 4734. of the Revised Code;

A Clinical nurse specialist licensed in Ohio pursuant to Chapter 4723. of the Revised Code can order x-ray procedures that are specified in the standards of care arrangement and under the scope of his or her collaborating physician; and

A physician assistant licensed in Ohio pursuant to Chapter 4730. of the Revised Code can order x-ray procedures that are specified in the utilization plan and under the scope of his or her collaborating physician

(18) An inventory of radiation-generating equipment, including the location and description of each unit.

The written quality assurance program must include an inventory of radiation-generating equipment.

Examples:

- There is a podiatric unit located in Room 8.

- There is an intraoral unit located in each of the following operatories: 1, 2/3, and 4.

(C) In addition to the obligations of paragraphs (A) and (B) of this rule, the quality assurance program of hospital registrants will comply with the following:

ONLY hospitals as defined in section 3727.01 of the Ohio Revised Code must follow paragraph (C) of this rule.

(1) A certified radiation expert will conduct oversight and maintenance of quality assurance programs for hospital registrants, by:

An Ohio Certified Radiation Expert (CRE) must conduct oversight and maintenance of the hospital’s quality assurance program for the following categories of radiation-generating equipment:

• 	Therapeutic

• 	Diagnostic other than mammography

• 	Mammography

A radiation expert is certified in accordance with rule 3701:1-66-03 of the Ohio administrative Code.

A hospital may use a single CRE certified in all applicable categories or the number of CREs necessary to cover all categories.

A current copy of the CRE’s certificate is required to verify the credentials of the CRE at inspections.

- Auditing the quality assurance program on an annual basis;

The CRE of the appropriate category must audit the hospital’s quality assurance program at intervals not to exceed one year, plus or minus one month (according to the definition of ‘annually’ found in rule 3701:1- 38-01(A)(18) of the OAC. The audit must be completed according to the requirements of rule 3701:1-66-04(C)(6) of the OAC.

- Performing reviews of the quality assurance program each quarter;

The CRE of the appropriate category must perform quarterly reviews in accordance with the requirements of rules 3701:1-66-04(C)(5)(a – e) of the OAC. The definition of ‘quarterly’ can be found in rule 3701:1- 38-01 of the OAC.

"Quarter" or "quarterly" means a period of time equal to one-fourth of the year observed by the licensee or registrant, approximately thirteen consecutive weeks, providing that the beginning of the first quarter in a year coincides with the starting date of the year and that no day is omitted or duplicated in consecutive quarters.

(c) Completing and submitting all necessary information with the annual audit form in accordance with paragraph (C)(6) of this rule; and

Using the Ohio Department of Health’s audit form, the CRE of the appropriate category is required to complete an annual audit of the hospital’s quality assurance program and submit all the required information and attachments to the hospital quality assurance committee and the Ohio Department of Health as required by paragraph (C)(6) of this rule.

(d) Serving on the quality assurance committee;

CRE(s) to cover the category of services (diagnostic, therapeutic, mammography) provided by the hospital must be on the quality assurance committee.

(2) Employees working in the radiation areas will be made aware of the identity, scope of authority, and a method for contacting the certified radiation expert and the individual responsible for radiation protection. This information, or a specific location where this information may be obtained, will be conspicuously posted in each area where radiation-generating equipment is used;

Documents containing the identity, scope of authority and method to contact the Certified Radiation Expert and Individual Responsible for Radiation Protection must be posted in each area where radiation-generating equipment is used. This can be a central location between x-rays rooms where routine work takes place. The posting is to be placed where individuals who work in the radiation area can observe the postings during their normal work routine.

(3) Each hospital registrant will establish a quality assurance committee for the management of the quality assurance program. The members of the quality assurance committee will be approved by an executive administrator. Committee meetings may be attended by the members or similarly qualified, designated alternates. The quality assurance committee will include at least the following members:

A committee for the management of the quality assurance program must be established by each hospital registrant. Members and alternate members on this committee must be approved by a hospital executive administrator. Attendance by either the committee member or the similarly qualified individual designated as their alternate may count towards meeting the quorum requirements described in paragraph (C)(4) of this rule.

“Similarly Qualified” is taken to mean having at least the same license or certification, (i.e. An alternate for the hospital CRE would be another CRE).

(a) A member of the hospital's executive administration;

An individual employed in the hospital’s administration must be a member of the quality assurance committee. According to the definition of “executive administration” found in rule 3701:1-66-01 of the Ohio Administrative code, this individual must have the authority to expend capital funds, approve personnel actions and implement changes to hospital policy and procedure.

(b) The individual responsible for radiation protection;

The individual designated by the registrant on their registration must be a member of the quality assurance committee. The Individual Responsible for Radiation Protection (IRRP) must meet the qualification of Ohio Administrative Code 3701:1-66-02(C)(3).

(c) A radiologist or radiation oncologist;

The quality assurance committee must have a member who meets one of the following:

An Ohio physician board certified by “the American Board of Radiology,” “the American Osteopathic Board of Radiology” or “the Royal College of Physicians and Surgeons of Canada;” or

- Who has completed a formal (residency) training program approved by the “Radiology Residency Review Committee of the Accreditation Council for Graduate Medical Education;” or

- Who has completed a formal (residency) training program approved by the “Education Council on Postdoctoral Programs of the American Osteopathic Association;” or

An Ohio physician certified in one of the following:

Radiation oncology or therapeutic radiology by the "American Board of Radiology" or combined diagnostic and therapeutic radiology program by the "American Board of Radiology" prior to 1976; or

- Radiation oncology by the "American Osteopathic Board of Radiology"; or

- Radiology, with specialization in radiotherapy, as a British "Fellow of the Faculty of Radiology" or "Fellow of the Royal College of Radiology"; or

- Therapeutic radiology by the "Canadian Royal College of Physicians and Surgeons.

(d) A certified radiation expert representing each of the following as applicable in each hospital;

(i) Radiation therapy services,

(ii) Mammography, or

(iii) Diagnostic radiography other than mammography; and

The certified radiation expert(s) of the appropriate category must be a member of the quality assurance committee.

(e) A management representative of each department of the hospital which has responsibilities involving the handling of radiation-generating equipment;

A management representative from each department which has responsibilities for handling radiation-generating equipment must be included in the membership of the quality assurance committee.

A member of management is any individual who is in a position of authority over any number of other individuals in the represented department of the hospital.

(4) The quality assurance committee will meet as often as is deemed necessary to carry out its duties, but at least annually. To establish a quorum at least one-half of the committee's membership will be present either in person or by telecommunication, and will include the individual responsible for radiation protection for the hospital, and the member of the executive administration of the hospital. A record of each meeting will be maintained and distributed to each member which will include the following:

At a minimum, the quality assurance committee must meet at least annually.

Although attendance by all members or designated alternates is encouraged, at a minimum, each member or his/her designated alternate must be present for at least one meeting each calendar year.

At least one-half of the committee’s membership must be present either in person or by telecommunication to establish a quorum. For a committee with an odd number of members, quorum is reached by rounding up (i.e. a 7 member committee would require attendance by 4 members to meet quorum).

All meetings require the attendance of the individual responsible for radiation protection and the executive administrator or their approved alternates.

A record of the meeting shall be maintained and a process in place to distribute the meeting record to each appointed member of the committee.

(a) The date of the meeting;

The record of the Quality Assurance Committee meeting must include the date of the meeting.

(b) An indication of members present; and

The record of the Quality Assurance Committee meeting must include the identification of the members who attended the meeting. This includes attendance by telecommunication.

(c) A summary of meeting including any recommended actions and ALARA reviews;

The record of the Quality Assurance Committee meeting must include a summary of the meeting and all recommended actions and ALARA reviews.

(5) Each quarter, the certified radiation expert will submit, to each appointed quality assurance committee member, a review of the quality assurance program, which will contain, as applicable:

Quarter or quarterly means a period of time equal to one-fourth of the year observed by the licensee or registrant, approximately thirteen consecutive weeks, providing that the beginning of the first quarter in a year coincides with the starting date of the year and that no day is omitted or duplicated in consecutive quarters.

The certified radiation expert of the appropriate category (Diagnostic, therapeutic, mammography) shall submit, to each member of the quality assurance committee, a review of the hospital quality assurance program which shall contain, as applicable.

Inspectors will verify that each member of the quality assurance committee received the quarterly review.

(a) Radiation safety policy revisions proposed by the certified radiation expert;

If the certified radiation expert proposes any revisions to the radiation safety policies of the quality assurance program during the quarter, the policy revisions must be submitted as part of the quarterly review to each quality assurance committee member.

If the CRE does not propose any revisions, it should be documented that no revisions are proposed in the review submitted to the quality assurance committee members.

(b) A review of occupational exposure records by the certified radiation expert;

The certified radiation expert must submit to the quality assurance committee members a review of the occupational exposure records as part of the quarterly review of the quality assurance program.

(c) Radiation safety incidents;

The certified radiation expert must submit to the quality assurance committee members any radiation safety incidents identified during the quarterly review.

If there are no radiation safety incidents, it should be documented that no radiation safety incidents occurred on the review submitted to the quality assurance committee members.

(d) Performance evaluation summaries for radiation-generating equipment including a description of any issues found; and

A summary of the radiation-equipment performance evaluations that took place during the reviewed quarter must be included in the certified radiation expert’s review submitted to the quality assurance committee. The summary must include whether issues are found and if an issue is found a description of the issue found.

(e) Any corrective actions recommended by the certified radiation expert that are necessary to comply with the obligations of this chapter;

If the certified radiation expert identifies any areas of noncompliance with the regulations, the recommended corrective actions to the non-compliance must be submitted to the quality assurance committed as part of the quarterly review of the quality assurance program.

If the certified radiation expert does not identify non-compliances with the regulations, it should be documented in the review submitted to the quality assurance committed.

(6) The quality assurance program will be audited at least annually by a certified radiation expert. The certified radiation expert will develop a written report of the audit findings on forms prescribed by the director and submit the report to the quality assurance committee within thirty days of completing the audit. The quality assurance committee will review the audit report and implement any corrective actions determined to be necessary. The certified radiation expert will file the audit report with the director within ninety days of completing the audit. Every audit report will include a determination of whether the quality assurance program properly addresses the matters described in this rule and whether it is being carried out in accordance with the written quality assurance program, and any corrective actions to be taken to comply with the obligations of this chapter. The audit report will become a part of the inspection record.

A certified radiation expert, certified in each category of radiation-generating equipment handled by the hospital, must audit the hospital’s quality assurance program annually.  Annually is defined in Ohio Administrative Code 3701:1-38-01.

The certified radiation expert’s written audit findings must be on the audit form provided by the Ohio Department of Health and include the CRE’s:

- Determination of whether the quality assurance program properly addresses all requirements of rule 3701:1-66-04 of the OAC and whether the quality assurance program is being carried out as written.

- Determination of any corrective actions to be taken to comply with the requirements of Chapter 3701:1-66 of the OAC.

The certified radiation expert must submit the annual audit report to each member of the quality assurance committee within thirty days of completing the audit.

The quality assurance committee must review the annual audit report and implement any corrective actions determined to be necessary. The committee’s review and corrective action must be documented to demonstrate occurrence. A good place for this documentation is in the quality assurance meeting minutes.

The certified radiation expert must file the audit report with the Ohio Department of Health within 90 days of completing the audit.

(D) In addition to the obligations of paragraphs (A) and (B) of this rule, the quality assurance program of registrants performing fluoroscopically-guided interventional other than veterinary procedures, and computed tomography (CT) other than veterinary and cone beam CT procedures will establish a radiation dose review committee in accordance with the following:

Registrants performing fluoroscopically-guided interventional and computed tomography procedures must establish a radiation dose review committee.

Veterinary fluoroscopically-guided interventional and computed tomography use are exempt from this requirement.

Cone Beam Computed Tomography use is exempt from this requirement.

(1) The registrant may establish a system-wide committee if the registrant has more than one site;

If the registrant has more than one site, the registrant may establish a system wide committee to address all facilities on the registration.  Registration of more than one site is determined using Ohio Administrative Code 3701:1-38-03(B).

(2) If the registrant is a subsidiary of a hospital, the obligations of paragraph (D) of this rule may be delegated to the hospital quality assurance committee provided its members meet the obligations of paragraph (D)(3) of this rule;

If the registrant is owned or controlled by a hospital, the hospital’s quality assurance committee can perform the radiation dose review committee’s function for that registrant provided the quality assurance committee members meet the following:

- Individuals Responsible for Radiation Protection (IRRP);

- Diagnostic Radiation Expert;

- As applicable, a physician that performs fluoroscopically-guided interventional and/or computed tomography procedures; and

- As applicable, a technologist that performs fluoroscopically-guided interventional and/or computed tomography procedures.

(3) The radiation dose review committee will include at least the following members:

(a) The individual responsible for radiation protection;

The individual designated by the registrant on the registration must be a member of the radiation dose review committee.

(b) A diagnostic radiation expert;

A diagnostic radiation expert must be a member of the radiation dose review committee.

(c) As applicable, a physician that performs fluoroscopically-guided interventional and/or computed tomography procedures; and

If fluoroscopically-guided interventional procedures are performed, a physician that performs the fluoroscopically-guided interventional procedures must be a member of the radiation dose review committee.

If computed tomography procedures are performed, a physician that performs computed tomography procedures must be a member of the radiation dose review committee.

(d) As applicable, a technologist that performs fluoroscopically-guided interventional and/or computed tomography procedures;

If fluoroscopically-guided interventional procedures are performed, a technologist that performs fluoroscopically-guided interventional procedures must be a member of the radiation dose review committee.

If computed tomography procedures are performed, a technologist that performs computed tomography procedures must be a member of the radiation dose review committee.

(4) A quorum of the radiation dose review committee will meet as often as necessary to carry out its duties, but at least annually. To establish a quorum at least one-half of the committee's membership will be present either in person or by telecommunication, and will include the individual responsible for radiation protection. A record of each meeting will be maintained and include the following:

At a minimum, the quality assurance committee members must meet at least annually.

Annually means at intervals not to exceed one year, plus or minus one month.

At least one-half of the committee members and the individual responsible for radiation protection must be present either in person or by telecommunication to establish a quorum for the meeting.

A record of the meeting shall be maintained.

(a) The date of the meeting;

The record of the radiation dose review committee meeting must include the date of the meeting.

(b) An indication of members present; and

The record of the radiation dose review committee meeting must identify the members present. This includes attendance by telecommunication.

(c) A summary of meeting including any recommended actions;

The record of the radiation dose review committee meeting must include a summary of the meeting and any recommended actions.

(5) The radiation dose review committee for fluoroscopically-guided interventional procedures will establish and implement written policies that include but are not limited to the following:

(a) Identification of individuals who are authorized to use fluoroscopic systems for interventional purpose;

The radiation dose review committee for fluoroscopically-guided interventional procedures must develop a list identifying authorized users.

(b) A method to be used to monitor patient radiation dose during fluoroscopically-guided interventional procedures;

The radiation dose review committee for fluoroscopically-guided interventional procedures must develop policies addressing methods to monitor patient dose.

(c) Dose notification levels, as appropriate, at which the physician is notified and appropriate actions are taken for patient safety;

The radiation dose review committee for fluoroscopically-guided interventional procedures must develop policies addressing dose notification levels used to notify the physician to determine appropriate action for patient safety.

(d) Substantial radiation dose level values following nationally recognized standards;

The radiation dose review committee for fluoroscopically-guided interventional procedures must develop policies addressing substantial radiation dose levels following nationally recognized standards.

(e) Actions to be taken for cases when a substantial radiation dose level is exceeded which may include patient follow-up; and

The radiation dose review committee for fluoroscopically-guided interventional procedures must develop policies addressing actions to be taken if a substantial dose level is exceeded.

(f) Reviewing policies identified in paragraphs (C)(5)(a) to (C)(5)(e) of this rule at least annually;

The radiation dose review committee for fluoroscopically-guided interventional procedures must review these policies required by Ohio Administrative Code 3701:1-66-04(D)(5) at least annually.

Annually means at intervals not to exceed one year, plus or minus one month.

(6) The radiation dose review committee for computed tomography will determine and review written protocols to improve image quality and minimize patient dose. The review will include acquisition and reconstruction protocols, image quality, and radiation dose. At a minimum, the review will be performed annually and include the following clinical protocols, if performed:

The dose review committee for Computed Tomography (CT) must determine and review written protocols to improve image quality and minimize patient dose.

The CT dose review committee’s annual review of the written protocols must include image acquisition and reconstruction, image quality and radiation dose.

At a minimum, the review of the written protocols must include the following clinical protocols if performed:

(a) Pediatric head;

(b) Pediatric abdomen;

(c) Adult head;

(d) Adult abdomen;

(e) Adult chest; and

(f) Brain perfusion.

(E) Records obligated by this chapter and Chapter 3701:1-67 of the Administrative Code will be maintained in accordance with the following:

(1) Data and test results of evaluations and calibrations of all radiation-generating equipment for no less than five years;

The records for data and test results of evaluations and calibrations of radiation-generating equipment must be kept for no less than five years.  Data and test results may be numerical such as: kVp setting, timer setting, reproducibility, dose rates, assumptions, etc.

This includes all the data from the evaluations of newly installed units and the accompanying data of service records when units are found operating outside the applicable limits specified in the Ohio Administrative Code.

(2) Data and test results of evaluations of shielding and surroundings of all radiation-generating equipment until the director terminates the registration or five years after the equipment is transferred or disposed;

The records for data and test results of evaluations of shielding and surroundings (area radiation surveys) must be kept until the director terminates the registration or five years after the equipment is transferred or disposed. Data and test results may be numerical such as: kVp setting, timer setting, dose rates, assumptions, etc.

(3) Maintenance logs for radiation-generating equipment for five years;

The maintenance logs of radiation-generating equipment must be kept for five years.

(4) Medical event reports involving radiation exposure to individuals for all radiation-generating equipment until the director terminates the registration;

All medical event reports must be maintained until the registration is terminated.

(5) Copies of current licenses or the department's licensure verification web page for everyone who is obligated to possess a license at the facility; and

A copy of the current license or a copy of the Ohio Department of Health’s license verification web page for each radiation-generating equipment operator or nuclear medicine technologist must be available at the facility.

(6) Biennial calibration certificates or cross calibration documentation for all instruments used to perform area radiation surveys, calibrations, and evaluations for five years.

Ensure companies hired to perform these services provide the calibration certificates of their testing equipment used to perform the services. Often a company will put the information directly on the maintenance document.