3701:1-66-02   General administration obligations for medical radiation-generating equipment.

As used in this rule, "radiation-generating equipment" means radiation-generating equipment used for dental, veterinary, or medical purpose but does not include therapeutic radiation-generating equipment.

- The director may, upon application thereof or upon his or her own initiative, grant a variance to the obligations of rules in this chapter as he or she determines is authorized by law, provided that the registrant shows to the satisfaction of the director that there is good cause for the variance, and that the variance will not result in any undue hazard or effect on the public health and safety or environment. The terms, conditions, and expiration of the variance will be set forth in writing by the director. Failure to comply with the terms of the variance may result in immediate revocation of the variance.

A request for a variance must be made in writing to the Director and must show good reason for the variance and that the variance will not result in any undue hazard or effect on public health and safety or environment.

Variances that are approved by the Director will set the terms, conditions and expiration of the variance. If the terms and conditions of the variance are not met, the variance may be revoked.

If the variance has expired or the effective date of the rule in which the variance was granted has changed, the variance is invalid, and a new variance must be requested.

A copy of the approved variance must be maintained on site.

If a variance has been granted, it is best to provide the inspector with a copy of the variance before the inspection begins. The inspector will verify the expiration date and that the terms outlined in the variance are being met.

- An individual will not expose any individual or animal to the useful beam unless ordered by a licensed practitioner acting within his or her scope-of-practice for dental, medical or radiation therapy purposes. Exposing an individual for training, demonstration or other purposes is forbidden unless otherwise specified in rules promulgated under Chapter 4773 or 3748 of the Revised Code.

Rule 3701:1-66-17 of the Administrative Code addresses medical screening and research. Exposing individuals for training, demonstration, or other non-dental or non-medical diagnostic purposes is prohibited.

Only a patient is to be exposed to the useful beam and the exposure must be authorized by licensed practitioner. A licensed practitioner is one of the following:

- A Dentist licensed in Ohio pursuant to Chapter 4715. of the Revised Code;

- A Physician licensed in Ohio to practice medicine or surgery or osteopathic medicine or surgery pursuant to Chapter 4731. of the Revised Code;

- A Podiatrist licensed in Ohio pursuant to Chapter 4731. of the Revised Code;

- A Veterinarian licensed in Ohio pursuant to Chapter 4741. of the Revised Code, can order x-ray on animals only;

- A Chiropractor licensed in Ohio pursuant to Chapter 4734. of the Revised Code;

- A Clinical nurse specialist licensed in Ohio pursuant to Chapter 4723. of the Revised Code can order x-ray procedures that are specified in the standards of care arrangement and under the scope of his or her collaborating physician; and

- A physician assistant licensed in Ohio pursuant to Chapter 4730. of the Revised Code can order x-ray procedures that are specified in the utilization plan and under the scope of his or her collaborating physician.

(C) The handler will assure:

(1) Every individual who performs radiologic procedures on human beings holds the appropriate radiologic license as obligated by Chapter 3701-72 of the Administrative Code and Chapter 4715. of the Revised Code.

The handler shall assure that every individual who performs radiologic procedures on human beings holds the appropriate radiologic license as required by Ohio Administrative Code 3701-72 and Ohio Revised Code 4715.

License categories:

- Radiographers

- General X-ray Machine Operator

- Nuclear Medicine Technologist

- Radiation Therapy Technologist or Radiation Therapist

(2) Every individual who is licensed to perform radiologic procedures is adequately instructed in the registrant's safe operating procedures and can demonstrate competency in the safe use of the equipment.

Selection of techniques, proper positioning, and knowledge of safe operating procedures are some of the items to be demonstrated.

Documented sign-off on the safe operating procedures or using a checkoff list showing the date of the competency checks and the signature of the operator and person conducting the competency test is acceptable.

(3) The individual responsible for radiation protection (IRRP) is qualified as one of the following:

Inspectors will verify that the IRRP is qualified in one of the categories listed below.

- Ohio licensed to operate radiation-generating equipment;

A radiographer, radiation therapist, or general x-ray machine operator licensed in Ohio pursuant to Chapter 4773. of the Revised Code.

A Dentist licensed in Ohio pursuant to Chapter 4715. of the Revised Code;

A Dental Hygienist licensed in Ohio pursuant to Chapter 4715. of the Revised Code;

A Physician licensed in Ohio to practice medicine or surgery or osteopathic medicine or surgery pursuant to Chapter 4731. of the Revised Code;

A Podiatrist licensed in Ohio pursuant to Chapter 4731. of the Revised Code;

A Veterinarian licensed in Ohio pursuant to Chapter 4741. of the Revised Code for animals only;

A Chiropractor licensed in Ohio pursuant to Chapter 4734. of the Revised Code;

- Dental assistant certified to operate dental radiation-generating equipment;

A dental assistant licensed in Ohio pursuant to Chapter 4715. of the Revised Code.

- Registered veterinary technician and trained to operate veterinary radiation-generating equipment;

A veterinary technician registered in Ohio pursuant to Chapter 4741. of the Revised Code.

- Certified by the American registry of radiologic technologists in a pathway involving ionizing radiation or certified by the nuclear medicine technologist certification board;

Radiographers, radiation therapists, nuclear medicine technologists and limited scope radiographers certified by the American Registry of Radiologic Technologists or Nuclear Medicine Technologist Certification Board.

- A radiation expert as defined in rule 3701:1-66-01 of the Administrative Code;

"Radiation expert" means an individual who meets the qualifications of:

- Applicable paragraphs of rule 3701:1-66-03 of the Administrative Code;

- Paragraph (D) of rule 3701-83-45 of the Administrative Code, for any facility providing radiation therapy services;

- Paragraph (C)(3) of rule 3701-83-52 of the Administrative Code for CT equipment, or paragraph (F)(3) of rule 3701-83-52 of the Administrative Code for fluoroscopy, at any facility providing CT or fluoroscopy services; or

- 21 C.F.R. 900.12(a)(3) (as effective on the effective date of this rule) for any facility providing mammography services.

- A health physicist certified by the American board of health physics; or

A health physicist certified by the American Board of Health Physics.

- An associate's degree or higher in health physics, radiologic science, nuclear medicine or nuclear engineering.

An individual with an associate’s degree or higher in health physics, radiologic science, nuclear medicine or nuclear engineering.

- Any radiation-generating equipment that does not meet the provisions set forth in this rule or any other applicable equipment obligations of Chapter 3701:1-66 of the Administrative Code will not be used to irradiate patients unless the director or a radiation expert determines that the non-compliance will not pose a radiation risk and arrangements have been made to promptly correct the non-compliance.

Each facility that is operating radiation-generating equipment for diagnostic purposes must meet the requirements of Chapter 3701:1-66, unless there is a variance from the director OR if a radiation expert has determined that the non-compliance will not pose a radiation risk or other safety hazard to workers or patients and arrangements have been made to correct the noncompliance.

"Radiation expert" means an individual who meets the qualifications of:

- Applicable paragraphs of rule 3701:1-66-03 of the Administrative Code;

- Paragraph (D) of rule 3701-83-45 of the Administrative Code, for any facility providing radiation therapy services;

- Paragraph (C)(3) of rule 3701-83-52 of the Administrative Code for CT equipment, or paragraph (F)(3) of rule 3701-83-52 of the Administrative Code for fluoroscopy, at any facility providing CT or fluoroscopy services; or

- 21 C.F.R. 900.12(a)(3) (as effective on the effective date of this rule) for any facility providing mammography services.

- Radiation-generating equipment will bear a warning label on the control panel, by the exposure switch or by the main power switch which cautions individuals that radiation is produced when it is energized.

All radiation-generating equipment must have a warning label on the control panel. The use of universal symbols is acceptable for meeting the requirements of this rule.

(F) Unless otherwise specified in this paragraph, radiation-generating equipment will meet the following standards:

(1) On battery-powered x-ray generators, visual means will be provided on the control panel to indicate whether the battery is in a state of charge adequate for proper operation;

All battery-powered generators must have an indicator on the control panel showing the amount of charge remaining on the battery for adequate operation, so that each operator knows the unit will operate properly during the x-ray exposure.

(2) The leakage radiation from the diagnostic source assembly measured at a distance of one meter in any direction from the source will not exceed 0.88 milligray air kerma (one hundred milliroentgen exposure) in one hour when the x-ray tube is operated at its leakage technique factors. Compliance will be determined by measurements averaged over an area of one hundred square centimeters with no linear dimension greater than twenty centimeters;

All FDA compliant equipment will meet this requirement. It is the responsibility of the manufacturer to ensure that leakage radiation shall not exceed one hundred milliroentgens in one hour.

(3) Except for mammographic radiation-generating equipment, the half-value layer (HVL) of the useful beam for a given x-ray tube potential will not be less than the values shown in table 1. If it is necessary to determine such HVL at an x-ray tube potential which is not listed in table 1, linear interpolation or extrapolation may be made;

A material’s half-Value Layer (HVL) is the thickness of the material that reduces the intensity of the radiation entering the material to one-half of its original intensity.

The HVL of the useful beam of radiation-generating equipment, other than mammography equipment, must not be less than the HVLs shown in table 1 below.

If it is necessary to determine HVL at an x-ray tube kilovoltage peak not listed in table 1, linear interpolation may be used.

Table 1.

| X-Ray Tube Voltage (kilovolt peak) |  | Minimum HVL (millimeter of aluminum) |  |  |
| --- | --- | --- | --- | --- |
| Designed Operating Range | Measured Operating Potential | Specified Dental Systems1 | I- Other X-Ray Systems2 | II - Other X-Ray Systems3 |
| Below 51 | 30 | 1.5 | 0.3 | 0.3 |
|  | 40 | 1.5 | 0.4 | 0.4 |
|  | 50 | 1.5 | 0.5 | 0.5 |
| 51 to 70 | 51 | 1.5 | 1.2 | 1.3 |
|  | 60 | 1.5 | 1.3 | 1.5 |
|  | 70 | 1.5 | 1.5 | 1.8 |
| Above 70 | 71 | 2.1 | 2.1 | 2.5 |
|  | 80 | 2.3 | 2.3 | 2.9 |
|  | 90 | 2.5 | 2.5 | 3.2 |
|  | 100 | 2.7 | 2.7 | 3.6 |
|  | 110 | 3.0 | 3.0 | 3.9 |
|  | 120 | 3.2 | 3.2 | 4.3 |
|  | 130 | 3.5 | 3.5 | 4.7 |
|  | 140 | 3.8 | 3.8 | 5.0 |
|  | 150 | 4.1 | 4.1 | 5.4 |
| 1 Dental x-ray systems designed for use with intraoral image receptors and manufactured after December 1, 1980. |  |  |  |  |
| 2 Dental x-ray systems designed for use with intraoral image receptors and manufactured before or on December 1, 1980, and all other x-ray systems subject to this section and manufactured before June 10, 2006. |  |  |  |  |
| 3 All x-ray systems, except dental x-ray systems designed for use with intraoral image receptors, subject to this section and manufactured on or after June 10, 2006. |  |  |  |  |

(a) For capacitor energy storage equipment, compliance with the obligations of this paragraph will be determined with the system fully charged and a setting of ten milliampere-seconds (mAs) for each exposure; and

(b) The obligated minimal HVL of the useful beam will include the filtration contributed by all materials which are permanently between the source and the patient;

"Half-value layer (HVL)" means the thickness of specified material which attenuates the beam of radiation to an extent such that the AKR is reduced by one-half of its original value.

When determining or evaluating the minimum half-Value Layer (HVL), the required HVL of the useful beam from radiation-generating equipment can include contributions from all the materials permanently between the radiation source of the equipment and the patient.

(4) For x-ray systems which have variable kilovolt peak (kVp) setting and variable filtration for the useful beam, a device will link the kVp selector with the filter and will prevent an exposure unless the minimum amount of filtration necessary to produce the HVL obligated by paragraph (F)(3) of this rule is in the useful beam for the given kVp which has been selected;

"Kilovoltage peak (kVp)" means the maximum value of the electrical potential difference between the cathode and the anode of the x-ray tube during an exposure.

FDA compliant equipment will meet this requirement. The HVL must meet the required specifications.

(5) Where two or more x-ray tubes are controlled by one exposure switch, the tube that has been selected will be clearly indicated prior to initiation of the exposure. This indication will be both on the x-ray control panel and for dental equipment at or near the selected tube housing assembly;

If two or more x-ray tubes are controlled by one exposure switch, make sure there is a means to determine which tube is selected. This should be clearly marked at the control panel and near each tube head assembly.

(6) The x-ray tube housing assembly supports will be adjusted such that the tube housing assembly will remain stable during an exposure unless tube housing movement is a designed function of the radiation-generating equipment;

When positioned for exposure, the tube housing assembly must remain in the position it was placed in and not drift during exposure.

(7) The technique factors to be used during an exposure will be indicated before the exposure begins. If automatic exposure controls are used, the technique factors which are set prior to the exposure will be indicated. This obligation may be met by permanent markings on equipment having fixed technique factors. Indication of technique factors will be visible from the operator's position except in the case of spot films taken during fluoroscopy procedures or dental intraoral or panoral films; and

All technique factors must be visible from the operator position. For automatic exposure controls, the technique factors set prior to exposure must be indicated. For equipment having fixed technique factors, permanent markings on the equipment are acceptable.

Exceptions: Spot films taken during fluoroscopy, dental intra-oral and dental panoral units.

Note: These exceptions do not mean that the technique factors are not required to be on the control panel. It only means that the technique factors do not need to be visible from operator position.

(8) All position locking, holding, and centering devices on radiation-generating equipment components will function as designed by the manufacturer.

All locking, holding, and centering devices on the radiation-generating equipment must function as designed by the manufacturer. Check these components routinely to make sure they are working properly and document the checks as part of the quality assurance program.

(G) In addition to other applicable radiation safety rules in Chapter 3701:1-66 of the Administrative Code, handlers of radiation-generating equipment will meet the following radiation safety obligations:

(1) Software-based technique selections, a chart, or a combination of the two will be provided in the vicinity of the radiation-generating equipment's control panel which specifies, for examinations performed with that system, the following information:

A software-based technique selection, a technique chart, or a combination of the two are required at the control of each x-ray unit.

- Patient's body part, radiographic projection, anatomical size or age, and the technique factors to be utilized for each;

The technique chart must specify the body part, radiographic projection, anatomical size or age and the technique to be used.

Example: Chest X-ray, Posterior-Anterior Projection, patient size 23cm, technique 120 kVp, 400 ma at 0.01 seconds

Example: Dental intraoral techniques should be posted for both adult and child settings.

- Type and size of the image receptor to be used;

Example: 14 X 17-inch film cassette

- Type and focal distance of the grid to be used, if any; and

Example: 72-inch focal distance grid

- Source-to-image receptor distance (SID) to be used, except for fluoroscopy, and dental intraoral or panoral radiography;

Example: 72-inch SID

(2) Except for patients who cannot be moved out of the room, only the staff, ancillary personnel or other persons needed for the medical procedure or training will be in the room during the radiologic procedure. Other than the patient being examined:

(a)  All individuals will be positioned such that no part of the body will be struck by the useful beam unless protected by not less than 0.5 millimeter lead equivalent material;

All individuals in the room for the medical procedure must be positioned so that they are clear of the useful beam unless protected with not less than 0.5-millimeter lead equivalent material.

(b)  The x-ray operator, other staff, ancillary personnel, and other persons needed for the medical procedure will be protected from the direct scatter radiation by protective aprons or whole body protective barriers of not less than 0.25 millimeter lead equivalent material; and

Any individuals required to be in the room for the medical procedure of a patient during the x-ray procedure of a patient must be protected from direct scatter radiation by protective aprons or whole-body protective barriers of at least 0.25-millimeter lead equivalent material.

(c)  Human patients who cannot be removed from the room will be protected from the direct scatter radiation by whole body protective barriers of not less than 0.25 millimeter lead equivalent material or will be so positioned that the nearest portion of the body is at least two meters (6.5 feet) from both the tube head and the nearest edge of the image receptor;

Any patients, who cannot be removed from the room during the exposure, must be protected from direct scatter radiation by whole body protective barriers or aprons of 0.25-millimeter lead equivalency, or must be at least two meters from both the tube head and image receptor.

This is meant for patients in emergency rooms or intensive care unit settings. Otherwise, all other patients should be removed from the room.

(3) If performing a radiologic procedure needs auxiliary support for holding a patient or an image receptor, the handler will ensure the following:

(a)  Mechanical holding devices will be used when the procedure permits their use in lieu of having an individual hold the patient or image receptor;

Mechanical holding devices shall be available and considered before using a person to hold the patient or image receptor.

(b)  Written safe operating procedures obligated by paragraph (B)(4) of rule 3701:1-66-04 of the Administrative Code will indicate the obligations for selecting someone to hold a patient or image receptor, and the procedure that will be followed. All individuals holding a patient or image receptor during radiation exposures will be at least eighteen years of age; and

If all resources have been exhausted and another individual must be used to assist in holding during a radiographic procedure, there must be written safe operating procedures for how the individual is selected for holding the patient. The procedure must indicate that the individual must be eighteen years of age and wear the appropriate lead equivalent aprons and gloves. The procedures shall indicate that no worker shall be dedicated as a routine holder of patients during radiation exposure. The procedures must not allow pregnant individuals to hold patients.

(c) No individual will routinely hold patients or image receptors during radiologic procedures;

No individual in the office shall be dedicated as a routine holder of patients during radiation exposure.

(4) The facility will have protective aprons and gloves available in sufficient numbers to provide protection to anyone who is involved with x-ray operations;

The facility shall have enough shielding devices to protect everyone involved with the x-ray procedures.

(5) Any radiation worker participating in fluoroscopic, veterinary, or mobile or portable x-ray procedures is obligated to wear an individual monitoring device unless the IRRP or radiation expert determines it is unlikely the radiation worker will receive in excess of the doses specified in paragraphs (B)(1)(a) to (B)(1)(c) of rule 3701:1-38-14 of the Administrative Code;

All radiation workers participating in fluoroscopic, veterinary, mobile, or portable x-ray procedures must wear individual monitoring devices.

If the IRRP or Radiation Expert has determined based on the workload and appropriate measurements and calculations that the radiation dose to an adult worker is unlikely to exceed 10% of 5rem annually, personnel badges can be discontinued. If the x-ray workload increases or geometry of the radiation source location changes in the room, then further evaluation or reinstatement of badges is required.

"Radiation expert" means an individual who meets the qualifications of:

- Applicable paragraphs of rule 3701:1-66-03 of the Administrative Code;

- Paragraph (D) of rule 3701-83-45 of the Administrative Code, for any facility providing radiation therapy services;

- Paragraph (C)(3) of rule 3701-83-52 of the Administrative Code for CT equipment, or paragraph (F)(3) of rule 3701-83-52 of the Administrative Code for fluoroscopy, at any facility providing CT or fluoroscopy services; or

- 21 C.F.R. 900.12(a)(3) (as effective on the effective date of this rule) for any facility providing mammography services.

(6) The entrance air kerma resulting from the technique used for the specified average adult human patient for routine diagnostic radiography will not exceed the values listed in table 2. The entrance air kerma resulting from the technique used for routine intraoral bitewing exams will not exceed the values listed in table 3. All values of entrance air kerma are specified as free-in-air, without backscatter. The corresponding entrance exposure in milliroentgens is listed in parentheses. Linear extrapolation or interpolation will be used for an x-ray tube potential (kVp) not listed in table 3;

The exposure techniques (kVp, mA x time, mAs) used for the adult human patient for routine diagnostic radiography shall not exceed the entrance air kerma values listed in table 2 below.

The exposure techniques (kVp, mA x time, mAs) used for the adult human patient for routine intraoral bitewing procedures shall not exceed the entrance are kerma values listed in table 3 below.

A high entrance air kerma value can result from film processing related problems, calibration issues, the use of expired film and inadequate half value layers.

| Table 2. |  |  |
| --- | --- | --- |
| Radiographic technique | Adult thickness cm | Entrance air kerma mGy (mR) |
| Chest (pa), (non-grid) | 23 | 0.26 (30) |
| Chest (pa), (grid) | 23 | 0.35 (40) |
| Abdomen (kub) | 23 | 5.26 (600) |
| Lumbo-sacral spine (ap) | 23 | 6.13 (700) |
| Thoracic spine (ap) | 23 | 3.50 (400) |
| Full spine | 23 | 3.50 (400) |
| Cervical spine (ap) | 13 | 1.75 (200) |
| Skull (lateral) | 15 | 1.75 (200) |
| Foot (dp) | 8 | 0.88 (100) |

| Table 3. |  |  |
| --- | --- | --- |
| Tube Voltage kVp | D-Speed Film mGy (mR) | F-Speed Film Digital Receptor mGy (mR) |
| 50 | 4.82 (550) | 2.45 (280) |
| 55 | 4.56 (520) | 2.19 (250) |
| 60 | 4.12 (470) | 1.93 (220) |
| 65 | 3.64 (415) | 1.66 (190) |
| 70 | 3.15 (360) | 1.45 (165) |
| 75 | 2.72 (310) | 1.23 (140) |
| 80 | 2.28 (260) | 1.01 (115) |
| 85 | 2.06 (235) | 0.92 (105) |
| 90 | 1.84 (210) | 0.83 (95) |
| 95 | 1.71 (195) | 0.74 (85) |
| 100 | 1.58 (180) | 0.61 (70) |

(7) Procedures and auxiliary equipment designed to minimize patient and radiation worker exposure will be utilized as follows:

(a) For facilities utilizing radiographic film, the speed of the screen and film combinations used will be the fastest speed consistent with the diagnostic objective of the examinations. Film cassettes without intensifying screens will not be used for any routine diagnostic radiography, with the exception of veterinary and specimen radiography;

The speed of the screen film combination shall be the fastest combination for the diagnostic objective. Film/screen manufactures should be able to help regarding their products

Diagnostic radiography procedures on human beings shall not be performed without intensifying screens.

Exceptions: Veterinary or specimen radiography

(b) Radiation-generating equipment subject to rule 3701:1-66-05 of the Administrative Code will not be utilized in procedures where the source-to-skin distance (SSD) is less than thirty centimeters, except for veterinary x-ray systems;

The Source-to-Skin (SSD) is not to be less than thirty centimeters for human use.

Veterinary units are exempt from this rule.

(c) If grids are used between the patient and the image receptor to decrease scatter to the image receptor and improve contrast, the grid will be:

(i) Properly aligned, with the x-ray tube side facing the correct direction, and the grid centered to the central ray; and

When using a grid between the patient and image receptor, the grid shall be facing the x-ray tube and properly aligned and centered to the x-ray tube.

(ii) The proper focal distance for the SID being used;

When using a grid between the patient and image receptor the proper source-to-image distance for the grid must be used.

(8) Except for radiation-generating equipment used for veterinary, portable, dental panoral, dental intraoral, lithotripsy, or bone densitometry applications, the operator will be behind a protective barrier, either in a separate room, in a protected booth, or behind a shield and be able to see the patient without leaving the protected barrier; and

The operator shall stand behind a protective barrier during the x-ray exposure and be able to view the patient from behind the protective barrier during the exposure without leaving the barrier.

Veterinary, portable, dental panoral, dental intraoral, lithotripsy, and bone densitometry are exempt from this rule.

(9) Each radiographic image, or a record linked with each radiographic image, will contain the following:

(a) Patient identification;

(b) Date of examination; and

(c)  Operator identification.

The radiographic image or a record linked to the radiographic image shall contain the patient identity, examination date and operator identity.

Examples: labels, computer-based identification, markers

(H) In addition to other applicable structural shielding obligations in Chapter 3701:1-66 of the Administrative Code, handlers of radiation-generating equipment will:

(1) For all units, except those used for bone densitometry, mammography, dental panoral or dental intraoral radiography:

(a) Use a radiation expert to prepare a shielding design to include specifications for all structural radiation barriers:

"Radiation expert" means an individual who meets the qualifications of:

- Applicable paragraphs of rule 3701:1-66-03 of the Administrative Code;

- Paragraph (D) of rule 3701-83-45 of the Administrative Code, for any facility providing radiation therapy services;

- Paragraph (C)(3) of rule 3701-83-52 of the Administrative Code for CT equipment, or paragraph (F)(3) of rule 3701-83-52 of the Administrative Code for fluoroscopy, at any facility providing CT or fluoroscopy services; or

- 21 C.F.R. 900.12(a)(3) (as effective on the effective date of this rule) for any facility providing mammography services.

(i)  Prior to new construction, or renovation; and

A radiation expert shall prepare a shielding design prior to new construction or renovation for all radiation-generating equipment except for bone densitometry, mammography, dental panoral or dental intraoral.

(ii)  For new radiation-generating equipment installations which might cause a significant increase in radiation hazard.

A radiation expert shall prepare a shielding design for new radiation-generating equipment installations which might cause a significant increase in radiation hazard.

Increase in radiation hazard from a new installation includes, change in geometrical location of unit or increased tube energy.

Bone densitometry, mammography, dental panoral or dental intraoral equipment are exempt from this rule.

Rule 3701:1-66-05(D) exempts mobile or portable radiation generating equipment from this requirement (unless it is used in a stationary capacity).

Rule 3701:1-66-05(F) exempts mobile and portable veterinary units from this requirement.

Rule 3701:1-66-05(G) exempts hand-held veterinary units from this requirement.

(b) Prior to patient use, use a radiation expert to determine compliance with exposure levels in accordance with rule 3701:1-38-14 of the Administrative Code by performing:

"Radiation expert" means an individual who meets the qualifications of:

- Applicable paragraphs of rule 3701:1-66-03 of the Administrative Code;

- Paragraph (D) of rule 3701-83-45 of the Administrative Code, for any facility providing radiation therapy services;

- Paragraph (C)(3) of rule 3701-83-52 of the Administrative Code for CT equipment, or paragraph (F)(3) of rule 3701-83-52 of the Administrative Code for fluoroscopy, at any facility providing CT or fluoroscopy services; or

- 21 C.F.R. 900.12(a)(3) (as effective on the effective date of this rule) for any facility providing mammography services.

(i) An area radiation survey for new installation of radiation-generating equipment.

A radiation expert shall conduct an area radiation survey for new radiation-generating equipment installations prior to patient use.

Bone densitometry, mammography, dental panoral or dental intraoral equipment are exempt from this rule.

Special Note: According to rule 3701:1-66-01, of the Ohio Administrative Code, “Mobile radiation-generating equipment” is limited to equipment that is NOT used in a fixed location. As such, a radiation survey is required for any fluoroscopic equipment used in a fixed location, regardless of equipment design as required by 3701:1-66-05(D)(1)(a).

(ii) An area radiation survey for reinstallation or after any change in structural shielding unless, in the documented determination of a radiation expert, the reinstallation or change will not cause a significant increase in radiation hazard.

A radiation expert shall conduct an area radiation survey after any change in the radiation-generating equipment or its structural shielding which might cause a significant increase in radiation hazard.

Increases in radiation hazard from changes in radiation-generating equipment or structural shielding may occur with increased tube energy, changes in geometrical location of equipment or changed in wall material.

If a written determination is provided by a radiation expert identifying the reinstallation or change will not cause a significant increase in radiation hazard, a new radiation survey is not necessary. However, the written determination and the initial radiation survey must be kept indefinitely (life of room) for review by inspectors.

Bone densitometry, mammography, dental panoral or dental intraoral equipment are exempt from this rule.

(c) Use a radiation expert to perform a re-calculation of area radiation survey results after any increase in clinical workload that exceeds the assumptions used in the existing radiation survey.

If the operating parameters, location and geometry of the new equipment installation is the same as those of the old equipment and the old survey documentation exist, a new survey does not need to be completed, but the old survey documentation must be kept indefinitely (life of room) for review by inspectors.

(d)  Obtain a written report of the shielding design and the area radiation survey. A copy of the report will be made available to the department's inspector upon request.

A written report of the shielding design shall be obtained from the radiation expert and kept on file for the duration of the installation. A copy of the report must be made available to the inspector upon request.

SHIELDING PLAN DOCUMENT

Diagrams and Documentation of X-ray room and surroundings showing:

- Location of X-ray unit, protected area for the control console, and patient viewing system

- Room dimensions with window and door locations

- All Beam orientations to be used

- If applicable, location of wall cassettes

E. If applicable, location of Dark room and film storage areas

F. Adjacent location to the x-ray room (e.g. lounges, waiting rooms, exam rooms, offices, reception area, rest rooms, corridors etc…)

G. Location, type, and thickness of material required for each radiation barrier

H. Results of area radiation surveys performed outside each radiation barrier

- Documentation of survey instrument used to include model, serial number, and copy of calibration certificate

Variations of any set or combination of documents containing the information above may be acceptable.

(2) Assure that no individual operates or permits the operation of radiation-generating equipment unless structural shielding and protective barriers are used such that no person other than the patient being examined will receive a total effective dose equivalent in excess of the limits prescribed in rules 3701:1-38-12 and 3701:1-38-13 of the Administrative Code.

The handler shall ensure that all protective barriers are in place and provide proper shielding prior to individuals operating the equipment.

(3) Provide a protective barrier either in a separate room, in a protected booth, or use a mobile barrier that will intercept the useful beam and any direct scattered radiation.

The handler shall provide protective barriers for its operators.  If a mobile barrier is used the Safe Operating Procedures should address the placement of the shield.

(4) Provide a window of lead equivalency affording protection equal to that obligated by the adjacent barrier, a television monitoring system, or a mirror system large enough and so placed that the operator can see the patient without having to leave the protected area during exposure.

The handler shall provide patient visualization mechanisms such as a window with lead equivalency protection equal to that required by adjacent barriers, a television monitor, or a mirror system so that the operator can remain behind the protective barrier and see the patient during exposure.

(5) Assure the stationary CT and mobile CT radiation-generating equipment used in a fixed location provides for two-way aural communications between the human patient and operator.

The handler shall assure that stationary computed tomography equipment and mobile computed tomography equipment used in a fixed location has two-way aural communication between the operator and patient.

(I) Notwithstanding paragraph (H)(1)(b)(ii) of this rule, reinstallation of radiation-generating equipment of the same operating parameters, location and geometry does not obligate another area radiation survey as long as the previous documented area radiation survey is maintained and available for inspection.

Another area radiation survey is not required for reinstallation of radiation-generating equipment of the same operating parameters, location and geometry provided that the previous radiation-generating equipment’s area radiation survey is maintained and available for inspection.

- Same operating parameters – max radiation output (i.e. kVp and time/mA combination or mAs).

- Same installation location (i.e. placement, site).

- Same geometry (i.e. orientation, direction, alignment).

(J) In addition to all applicable rules in Chapter 3701:1-66 of the Administrative Code, handlers of radiation-generating equipment will meet the following quality assurance obligations:

(1) X-ray systems and associated components used on humans and certified pursuant to 21 C.F.R. part 1020 (as effective on the effective date of this rule) will be maintained in compliance with applicable requirements of that standard, any modifications to the original components or systems will comply with that standard, and handlers will maintain documentation of compliance between inspections;

All radiation-generating equipment certified pursuant to 21 C.F.R. part 1020 that is used to expose humans must be maintained in compliance with the C.F.R. certification standards. Any modification to the original components or systems must comply with the C.F.R. standards.

(2) The handler will maintain the following information for all radiation-generating equipment for inspection by the department:

(a) User's manuals;

The handler (facility) must keep the user manual(s) for the radiation-generating equipment.

(b) Records of surveys, calibrations, maintenance, and modifications performed on the radiation-generating equipment which will be maintained between inspections; and

The handler must keep the survey, calibration, maintenance, and modification records for the radiation-generating equipment between inspections. The records must include data if applicable.

(c) A copy of all correspondence with the department regarding each piece of radiation-generating equipment;

The handler must keep all correspondence they provide to the Department regarding their radiation-generating equipment.

(3) Unless otherwise specified in another rule in this chapter, each installation using a piece of radiation-generating equipment and using analog image receptors, such as radiographic film, will have available suitable equipment for handling and processing radiographic images in accordance with the following provisions:

(a) For manually processing film:

(i) Developer and fixer tanks will be constructed of mechanically rigid, corrosion resistant material; and

The developer and fixer tanks for manual film processing must be constructed of rigid and corrosion resistant material.

(ii) The temperature of solutions in the tanks will be maintained within the range of 15.6 to 26.7 degrees Celsius (sixty to eighty degrees Fahrenheit). Film will be developed in accordance with the time-temperature relationships recommended by the film manufacturer, or in absence of such recommendations, with the following time-temperature chart:

The films must be developed with the time-temperature combination recommended by the film manufacturer, or in absence of a film manufacturer recommendation, the films must be developed using the following time-temperature chart.

| Time-Temperature Chart |  |  |
| --- | --- | --- |
| Thermometer Reading (Degrees) |  | Minimum Developing Time (Minutes) |
| °C | °F |  |
| 26.7 | 80 | 2 |
| 26.1 | 79 | 2 |
| 25.6 | 78 | 2.5 |
| 25.0 | 77 | 2.5 |
| 24.4 | 76 | 3 |
| 23.9 | 75 | 3 |
| 23.3 | 74 | 3.5 |
| 22.8 | 73 | 3.2 |
| 22.2 | 72 | 4 |
| 21.7 | 71 | 4 |
| 21.1 | 70 | 4.5 |
| 20.6 | 69 | 4.5 |
| 20.0 | 68 | 5 |
| 19.4 | 67 | 5.5 |
| 18.9 | 66 | 5.5 |
| 18.3 | 65 | 6 |
| 17.8 | 64 | 6.5 |
| 17.2 | 63 | 7 |
| 16.7 | 62 | 8 |
| 16.1 | 61 | 8.5 |
| 15.6 | 60 | 9.5 |

(iii) Devices will be utilized which will indicate the actual temperature of the developer and signal the passage of a preset time appropriate to the developing time needed;

Devices must be used to indicate the temperature of the developer and the appropriate passage of time the film is to be in developer.

The timing device must provide a method to preset the development time and signal the passage at the end of a preset time.

(b) For automatic processors and other closed processing systems:

(i) Films will be developed in accordance with the time-temperature relationships recommended by the film manufacturer; in the absence of such recommendations, the film will be developed using the following chart:

The films must be developed with the time-temperature combination recommended by the film manufacturer, or in absence of a film manufacturer recommendation, the films must be developed using the following time-temperature chart.

| Developer Temperature (Degrees) |  | Minimum Immersion Timea/ |
| --- | --- | --- |
| °C | °F | Seconds |
| 35.5 | 96 | 19 |
| 35 | 95 | 20 |
| 34.5 | 94 | 21 |
| 34 | 93 | 22 |
| 33.5 | 92 | 23 |
| 33 | 91 | 24 |
| 32 | 90 | 25 |
| 31.5 | 89 | 26 |
| 31 | 88 | 27 |
| 30.5 | 87 | 28 |
| 30 | 86 | 29 |
| 29.5 | 85 | 30 |
| a/Immersion time only, no crossover time included. |  |  |

(ii) The specified developer temperature and immersion time will be posted in the darkroom, on the automatic processor, or be readily available to the operator;

The dark room must contain a film manufacturer time/temperature chart (or the chart contained in 66-02 (J)). A specified developer temperature and immersion time must be available (alternately, a ‘drop’ or ‘transport’ time is acceptable and usually found in the processor operator manual). The facility must document that they are processing according to the recommended time-temperature relationship. If the processor does not have a built-in temperature gauge the facility must provide documentation showing periodic temperature measurement (note: the facility should not use a glass thermometer to avoid mercury contamination).

(c) Processing deviations from the obligations listed above will be documented by the handler in such manner that the obligations of this rule are shown to be met or exceeded, such as with extended processing, and special rapid chemistry;

If the facility deviates from the temperature or immersion time, they must document the deviation and explain how this meets or exceeds time/temperature requirements.

(d) Film processing solutions will be prepared in accordance with the directions given by the film manufacturer, and will be maintained in strength by replenishment or renewal so that full development is accomplished within the time specified by the manufacturer; and

Check the chemical or processor manufacturer’s recommendations and adhere to these recommendations. Documentation of processor maintenance to include replenishment or renewal of developer must be documented.

Note: Improper preparation of film processing solutions can result in decreased life expectancy of clinical images due to retention of residual fixer (hypo); poor film quality including increased base + fog, decreased contrast and/or decreased speed; or high ESE readings.

(4) Pass boxes, if provided, will be so constructed as to exclude light from entering the darkroom when cassettes are placed in or removed from the boxes, and will incorporate adequate shielding from stray radiation to prevent exposure of undeveloped film;

Pass boxes must be constructed to exclude light from entering the dark room when cassettes are placed in or removed from the pass boxes.  The pass boxes must be shielded from stray radiation.

(5) The darkroom will be light tight and use proper safelighting such that any film which would produce an optical density between one and two when exposed in a cassette to x-radiation and then processed will:

(a) Not suffer an increase in optical density greater than 0.1 when exposed in the darkroom for two minutes with all safelights on; and

(b) Not suffer an increase in optical density greater than 0.05 for mammography when exposed to the darkroom for two minutes with all safelights on;

There should be no light coming from around the door or ceiling panels. Other items to consider are that the wattage of the bulb and type of light filter agree with the film manufacturer’s recommendations.

(6) Darkrooms typically used by more than one individual will provide a method to prevent accidental entry of light while undeveloped films are being handled or processed;

There must be a method to prevent accidental entry or a policy that prevents accidental entry into the darkroom when used by more than one individual. A rotating door system, sign on the door, or light outside the darkroom that comes on when the door is shut is acceptable.

(7) Film will be stored in a cool, dry place and will be protected from exposure to stray radiation. Film in open packages will be stored in a light-tight container. If used, daylight film handling boxes will preclude fogging of the film;

Film must be stored in a cool, dry place away from exposure to stray radiation. Film in open packages shall be stored in light-tight containers.  If daylight film handling boxes are used, the boxes must prevent fogging of the film

(8) Expired x-ray film will not be used for diagnostic radiographs;

Check film packages routinely to prevent the use of expired x-ray film.

(9) Cassettes, intensifying screens, and computed radiographic imaging plates will be:

(a) Cleaned according to manufacturer's specifications or an alternate frequency approved and documented by a radiation expert in the quality assurance program;

The facility’s film cassettes, intensifying screens and computed radiographic imaging plates must be cleaned according to manufacturer specifications or in accordance with the frequency approved by a radiation expert as documented in the quality assurance program.

A log with signatures and dates to coincide with the facility’s policy is acceptable.

"Radiation expert" means an individual who meets the qualifications of:

- Applicable paragraphs of rule 3701:1-66-03 of the Administrative Code;

- Paragraph (D) of rule 3701-83-45 of the Administrative Code, for any facility providing radiation therapy services;

- Paragraph (C)(3) of rule 3701-83-52 of the Administrative Code for CT equipment, or paragraph (F)(3) of rule 3701-83-52 of the Administrative Code for fluoroscopy, at any facility providing CT or fluoroscopy services; or

- 21 C.F.R. 900.12(a)(3) (as effective on the effective date of this rule) for any facility providing mammography services.

(b) Inspected for damage; and

The facility’s film cassettes, intensifying screens and computed radiographic imaging plates must be inspected for damage according to the facility’s quality assurance program.

A log with signatures and dates to coincide with the facility’s quality assurance program is acceptable.

(c) Replaced as necessary to assure radiographs of good diagnostic quality;

If damaged, film cassettes, intensifying screens and computed radiographic imaging plates must be replaced.

(10) For those registrants employing computed and digital radiography imaging systems, the following will apply:

(a) If the computed radiography reader is located in the same room as the radiation-generating equipment and it is not behind a protective barrier, x-ray exposures will not be made during processing;

When the computed radiography reader is in the same room as the radiation-generating equipment, no processing shall occur during radiation exposure. This should be identified in the safe operating portion of the quality assurance program.

(b) Computed radiography plates will be processed as soon as possible after exposure, not to exceed eight hours under any circumstances;

Process the computed radiography plates as soon as possible after exposure and never allow processing to exceed eight hours after exposure.

(c) Computed radiography plates will be adequately shielded from stray radiation. Registrants will develop a process that will ensure that computed radiography plates are used frequently enough or erased at least weekly so as to produce diagnostic quality images; and

Document and implement a process to properly store and erase computed radiography plates frequently enough to assure diagnostic quality images. Follow manufacturer specifications.

(d) Facilities other than dental, podiatric, and veterinary will complete at least annually image quality evaluations appropriate for the equipment as established by a radiation expert or system manufacturer;

Facilities shall complete annual image quality evaluations established by a radiation expert or the system manufacturer for computed and digital radiography imaging systems.

"Radiation expert" means an individual who meets the qualifications of:

- Applicable paragraphs of rule 3701:1-66-03 of the Administrative Code;

- Paragraph (D) of rule 3701-83-45 of the Administrative Code, for any facility providing radiation therapy services;

- Paragraph (C)(3) of rule 3701-83-52 of the Administrative Code for CT equipment, or paragraph (F)(3) of rule 3701-83-52 of the Administrative Code for fluoroscopy, at any facility providing CT or fluoroscopy services; or

- 21 C.F.R. 900.12(a)(3) (as effective on the effective date of this rule) for any facility providing mammography services.

(11) Annual evaluation of the integrity of all necessary protective apparel.

Facilities shall evaluate the integrity of lead aprons and lead gloves annually, as these apparel items are required by regulation.

‘Annually’ means at intervals not to exceed one year, plus or minus one month.

(K) Upon discovery of a medical event, the handler will:

Medical event - means one or more of the following criteria have occurred to a human patient:

- Unintended skin dose to the same area in a single procedure greater than 2 sievert (200 rem);

- Unintended dose other than skin dose in a single procedure greater than:

  - 0.5 sievert (50 rem) to any organ; or

  - 0.05 sievert (5 rem) effective dose equivalent;

- Wrong patient or wrong site for entire procedure when the resultant dose is:

  - Greater than 0.5 sievert (50 rem) to any organ; or

  - Effective dose equivalent greater than or equal to 0.05 sievert (5 rem).

Unintended Dose or Unintended Skin Dose - means a patient radiation dose resulting from an error or equipment malfunction during a procedure.

- Contact the department regarding the medical event within one business day;

Upon discovery of a medical event contact the Ohio Department of Health within one business day.

(2) Provide a written report, including the analysis of the medical event, by a radiation expert to the department within fifteen business days of the medical event. The written report will include:

(a) The handler or registrant's name;

(b) The name of the prescribing physician;

(c) A brief description of the event including the body site, dose delivered and any critical structures involved;

(d) Why the event occurred;

(e) The effect, if any, on the individual who received the medical event;

(f) Actions, if any, that have been taken, or are planned, to prevent recurrence; and

(g) Certification that the handler notified the individual, or the individual's responsible relative or guardian, and if not, why not.

When a medical event occurs, a written report, as described above, must be provided to the department within 15 days of discovering the medical event.

(3) Provide a clinical summary to the prescribing physician and patient within fifteen business days; and

When a medical event occurs, a written report, as described above, must be provided to the Ohio Department of Health within 15 days of the medical event.

(4) Maintain record of the medical event as part of the patient's permanent medical record.

Medical events records must be maintained in the patient’s file.

(L) The written report in paragraph (K)(2) of this rule will not contain the individual's name or any other information that could lead to the identification of the individual.

The report supplied to the Ohio Department of Health must not contain confidential patient information.