3701:1-66-11     Medical bone densitometry radiation-generating equipment.

- Handlers of bone densitometry equipment will assure that the equipment is certified by the manufacturer in accordance with the obligations of 21 C.F.R. part 1010.2, (as effective on the effective date of this rule), and maintained in accordance with the manufacturer's specifications.

Bone densitometry equipment means radiation-generating equipment used for the medical purpose of quantifying bone density and mineral content by x-ray measurements through the bone and adjacent tissues.

Certified Bone Densitometry equipment will have a label or tag permanently affixed to or inscribed on the unit that is readily accessible for view to indicate that the unit has been certified by the manufacture in accordance with the requirements of 21 CFR 1010.2

- In addition to other applicable radiation safety obligations in Chapters 3701:1-38 and 3701:1-66 of the Administrative Code, the operator of bone densitometry equipment will be positioned at least one meter (3.3 feet) from the primary beam or behind a protective barrier containing a minimum of 0.25 millimeter of lead equivalent materials.

Operators of bone densitometry units must be positioned at least one meter from the primary x-ray beam or be behind a protective barrier containing a minimum of 0.25 millimeter of lead equivalent material.

The protective barrier can be a mobile shield or protective apron.

In accordance with rule 3701:1-66-02(H)(4), when a protective barrier is used, the operator must be able to view the patient during exposure.

(C) In addition to other applicable quality assurance obligations in rules 3701:1-66-02 and 3701:1-66-04 of the Administrative Code, handlers of bone densitometry equipment will comply with the following:

(1) The quality assurance program will include quality control test records which will be maintained between inspections and will include the following:

(a) A list of the tests to be performed as specified by the manufacturer;

The quality assurance program must indicate the quality control tests to be performed as specified by the manufacturer

(b) The frequency of performance as specified by the manufacturer;

The written quality assurance program must include how often the quality control tests are to be performed as specified by the manufacturer.

(c) The acceptability limits for each test as specified by the manufacturer; and

The written quality assurance program must include the acceptable limits of each quality control test performed as specified by the manufacturer.

(d) A brief description of the procedures and test equipment to be used for each test;

The written quality assurance program must include a brief description of the procedures and test equipment to be used for each test.

(2) Maintain a readily available copy of the manufacturer's operating manual; and

The operating manual must be available for review. If a hard copy is not available, instruction for accessing the manual online must be documented in the quality assurance program.

(3) The operator will advise the patient that the bone densitometry examination is a type of x-ray procedure.

The written quality assurance program must clearly describe the method the facility uses to ensure the operator advises the patient that the bone densitometry examination is a type of x-ray procedure.

Operators shall inform the patient that the scan is a type of x-ray procedure. This can be done verbally or through a consent form.