3701:1-66-05     Medical general purpose radiographic equipment.

The obligations of this rule do not apply to radiation-generating equipment used for dental intraoral or panoral, mammography, bone densitometry, computed tomography, fluoroscopy or spot film imaging, and equipment used in radiation therapy.

(A)  General purpose radiographic equipment will meet the following equipment standards:

(1)  A means will be provided for limiting the x-ray beam to the image receptor and area of clinical interest;

An aperture, diaphragm, variable collimator or positive beam limiting (PBL) device shall be available on general-purpose radiographic equipment to limit the x-ray field to the image receptor and area of clinical interest.

(2)  For radiographic equipment having a variable x-ray field limitation device, the limitation device will have means for independent stepless adjustment of both the length and width of the x-ray field;

If the radiographic equipment has a variable x-ray field limitation device, the variable x-ray field collimators must be able to adjust seamlessly (continuous motion) to variable length and width sizes for collimating to the area of clinical interest or the image receptor size.

(3)  For radiographic equipment that employs a light field for visually defining the perimeter of the x-ray field, the light source will be functional and the total misalignment of the edges of the visually defined field with the respective edges of the x-ray field along either the length or width of the visually defined field will not exceed two per cent of the source-to-image distance (SID);

Radiation-generating equipment that uses a light field to define the borders of the x-ray field must have a functional light source and the total misalignment of the edges of visually defined light field with respect to edges of the length or width of the x-ray field must be within plus or minus two percent of the source-to-image distance.

For example: at a 40-inch SID, the misalignment of the x-ray field to the light field must be within 0.8 inches (40” x 0.02) along each axis of the x-ray field (i.e. 0.5 inches off on the left plus 0.5 inches off on the right = 1.0 inches along the length of the x-ray field which would not meet compliance).

(4)  In addition to the obligations of paragraphs (A)(1), (A)(2) and (A)(3) of this rule stationary radiographic equipment having a variable x-ray field limitation device will have:

(a)  A means provided to indicate when the axis of the x-ray beam is perpendicular to the plane of the image receptor;

A leveling device or other mechanism must be available on stationary radiation-generating equipment that indicates the tube head is parallel to the image receptor and therefore, the axis of the beam is perpendicular to the plane of the image receptor.

(b) A means provided to align the center of the x-ray field with respect to the center of the image receptor to within two per cent of the SID, when the x-ray beam is perpendicular to the plane of the image receptor;

When the x-ray beam is perpendicular to the plane of the image receptor, a lighted cross hair or other mechanism marking the center of the x-ray field with respect to the center of the image receptor must be available and the two centers must not vary by more than 2 percent of the source-to-image distance.

For example: at a 40-inch source-to-image distance, the centers must be aligned to within ±0.8 inches. Similarly, at a 72-inch source-to-image distance, the centers must be aligned to within approximately ±1.4 inches.

(c) A means to indicate the SID to within two per cent. If it is a fixed SID, the distance will be indicated with a permanent marking;

Markings on the unit or a measuring device on the collimator must indicate the source-to-image distance to within 2 percent of the distance.

For fixed source-to-image distances, the distance must be indicated by a permanent marking. The distance must be within 2 percent of the indicated distance.

For example: at a 40-inch source-to-image distance, the SID indicator must be accurate to within ±0.8 inches (between 39.2” & 40.8”). Similarly, a 72-inch SID indicator must be accurate to within approximately ±1.4 inches.

(d)  The beam limiting device indicate numerically the field size in the plane of the image receptor to which it is adjusted and be accurate within two per cent of the SID; and

The accuracy of the indicated x-ray field size on the beam limiting device must be within 2 percent of the specified source-to-image distance.

(e)  Compliance measurements made to discrete SID's and image receptor dimensions in common clinical use, or at any other specific dimensions at which the beam-limiting device or its associated diagnostic x-ray system is uniquely designed to operate;

Testing for compliance will be done using image receptor dimensions in common clinical use or image receptor dimensions the x-ray system is designed to operate.

The image receptor dimensions, and source-to-image distance scales are indicated on the collimation device of the x-ray equipment.

(5)  General purpose radiographic equipment designed with only one image receptor size at a fixed SID will have:

This type of radiographic equipment uses a single diaphragm at a specific source-to-image distance to limit the x-ray field to the size of the image receptor. Example: dedicated chest units.

(a)  A means to limit the field at the plane of the image receptor to dimensions no greater than those of the image receptor, and to align the center of the x-ray field with the center of the image receptor within two per cent of the SID; or

A collimation device must be available to limit the x-ray field to the size of the image receptor. Many times, this is a diaphragm collimator used with the correct source-to-image distance.

There must be a method for aligning the center of the x-ray field with respect to the center of the image receptor and the two centers must not vary by more than 2 percent of the source-to-image distance. For example, at a 72-inch source-to-image distance, the centers must be aligned to within approximately ±1.4 inches; or

(b)  A means to align the field such that the x-ray field at the plane of the image receptor will not extend beyond any edge of the image receptor;

A method for aligning the x-ray field with respect to the plane of image receptor such that the x-ray field is within all edges of the image receptor.

(6)  General purpose radiographic equipment designed with multiple removable or selectable fixed apertures will have:

(a) A means to limit the x-ray field in the plane of the receptor so that such field does not exceed each dimension of the image receptor by more than two per cent of the SID when the axis of the x-ray beam is perpendicular to the plane of the image receptor;

The collimation device must be able to limit the x-ray field to the plane of the image receptor such that the x-ray field does not exceed each dimension of the image receptor by more than two percent of the source-to-image distance when the x-ray beam is perpendicular to the plan of the image receptor.

For example: an 8 X 10-inch film at a 40-inch source-to-image distance. The actual field size for the 8-inch film dimension must be no greater than 8.8 inches of the and no greater than 10.8 inches for the 10-inch dimension.

(b) A means to align the center of the x-ray field with the center of the image receptor to within two per cent of the SID or for equipment uniquely designed where the beam axis is intended to be offset from the center of the image receptor, the x-ray field at the plane of the image receptor will not extend beyond the image receptor; and

There must be a method for aligning the center of the x-ray field with respect to the center of the image receptor and the two centers must not vary by more than 2 percent of the source-to-image distance.

(c)  Paragraphs (A)(6)(a) and (A)(6)(b) of this rule may be met with either:

(i)  An assortment of removable, fixed-aperture, beam limiting devices sufficient to meet the obligation for each combination of image receptor size and SID for which the unit is designed with each such device having clear and permanent markings to indicate the image receptor size and SID for which it is designed; or

Removable and fix-aperture beam limiting devices must be available for each combination of image receptor size and source-to-image distances being used with the x-ray equipment.

The beam limiting devices must have clear and permanent markings indicating the image receptor size and the specified source-to-image distance for which they were designed and to be used.

(ii)  A beam-limiting device having multiple fixed apertures sufficient to meet the obligation for each combination of image receptor size and SID for which the unit is designed. Permanent, clearly legible markings will indicate the image receptor size and SID for which each aperture is designed and will indicate which aperture is in position for use;

Multiple fixed apertures must be available for the combination of image receptor sizes and source-to-image distances being used with the x-ray equipment.

The beam limiting apertures must have clear and permanent markings indicating the image receptor size and the specified source-to-image distance for which they were designed and to be used.

(7)  If a positive beam limitation (PBL) device is used, it will meet the following additional obligations:

When positive beam limitation is used, the following additional requirements must be met:

(a)  The PBL will prevent the production of x-rays when any of the following conditions are met:

(i) Either the length or width of the x-ray field in the plane of the image receptor differs from the corresponding image receptor dimensions by more than three per cent of the SID, except as permitted by paragraph (A)(7)(c) of this rule;

The positive beam limitation device shall prevent the production of x-rays when the x-ray field in the plane of the image receptor differs from either the length or width of the image receptor size by more than 3 percent of the source-to-image distance unless the operator chooses to manually collimate down to a smaller beam size.

For example: at a 40-inch source-to-image distance, when an 8 X 10-inch film-cassette is inserted in the bucky tray, the radiation-generating equipment shall not produce x-rays if either the dimension is greater than 9.2 inches or less than 6.8 inches for the 8-inch dimension or 11.2 inches and 8.8 inches for the 10-inch dimension.

(ii) The sum of the length and width differences as stated in paragraph (A)(7)(a)(i) of this rule without regard to sign exceeds four per cent of the SID; or

The positive beam limitation device shall prevent the production of x-rays when the sum of the length and width of the x-ray field (regardless of sign) in the plane of the image receptor differs from the image receptor size by plus or minus 4 percent of the source-to-image distance.

For example, at a 40-inch source-to-image distance, when an 8 x 10 inch film- cassette is inserted in the bucky tray, a field size of 9.2 inches in the 8 inch dimension (+1.2”) combined with a field size of 9.4 inches in the 10 inch dimension (-0.6”) sums to |1.2| + |-0.6| = 1.8 inches (without regard to the sign), would exceed 4 percent (1.6”) of the 40-inch source-to-image distance.

(iii) The beam-limiting device is at an SID for which PBL is not designed for sizing;

When the positive beam limitation device is not at a designed source-to-image distance for sizing, it shall prevent the production of x-rays.

(b)  Compliance with paragraph (A)(7)(a) of this rule will be determined:

(i) When the equipment indicates that the beam axis is perpendicular to the plane of the image receptor; and

The positive beam limiting device compliance testing of paragraph (A)(7)(a) of this rule must be determined when the equipment indicates the axis of the beam is perpendicular to the plane of the image receptor.

(ii) No sooner than five seconds after insertion of the image receptor;

The positive beam limiting device compliance testing of paragraph (A)(7)(a) of this rule must be delayed for at least five seconds after inserting the image receptor.

(c) The PBL system will be capable of operation, at the discretion of the operator, such that the size of the field may be made smaller than the size of the image receptor through stepless adjustment of the field size. The minimum field size at a SID of one hundred centimeters will be equal to or less than five centimeters by five centimeters;

The positive beam limiting device must allow the operator to manually collimate the x-ray field to a smaller size than the image receptor. The x-ray field collimator must be able to adjust seamlessly (continuous motion) to variable sizes smaller than the image receptor. The minimum field size at a SID of one hundred centimeters shall be equal to or less than five centimeters by five centimeters.

Note: when applicable, collimating to the body part of interest rather than the size of the image receptor is optimal.  This improves the quality of the image and reduces exposure from scatter radiation.

(d)  The PBL system will be designed such that if a change in image receptor does not cause an automatic return to PBL function as described in paragraph (A)(7)(a) of this rule, then any change of image receptor size or SID will cause an automatic return; and

Any change in image receptor size or source-to-image distance shall cause an automatic return to positive beam limitation function.

(e)  The PBL system will function as described in paragraph (A)(7) of this rule whenever all the following conditions are met:

Whenever all the conditions described below are met, the positive beam limiting system shall function as described in paragraph (A)(7), which is summarized below:

(i) The image receptor is inserted into a permanently mounted cassette holder;

When the image receptor is inserted into a permanently mounted cassette holder:

- The positive beam limiting device must prevent the production of x-rays if:

  - Either the length or the width of the x-ray field differs from the image receptor size by three percent of the source-to-image distance.

  - The sum of the differences between the length and the width of the x-ray fields versus the image receptor size exceeds four percent.

- The positive beam limiting device must prevent the production of x-rays when not at a designed source-to-image distance for sizing.

- The positive beam limiting device must allow the operator to manually collimate the x-ray field to a smaller size than the image receptor.

- A change in image receptor size must cause an automatic return to positive beam limitation function.

(ii) The image receptor length and width are less than fifty centimeters;

When the image receptor length and width are less than fifty centimeters:

- The positive beam limiting device must prevent the production of x-rays if:

  - Either the length or the width of the x-ray field differs from the image receptor size by three percent of the source-to-image distance.

  - The sum of the differences between the length and the width of the x-ray fields versus the image receptor size exceeds four percent.

- The positive beam limiting device must prevent the production of x-rays when not at a designed source-to-image distance for sizing.

- The positive beam limiting device must allow the operator to manually collimate the x-ray field to a smaller size than the image receptor.

- A change in image receptor size must cause an automatic return to positive beam limitation function.

(iii) The x-ray beam axis is within plus or minus three degrees of vertical in any direction and the SID is ninety to one hundred thirty centimeters inclusive; or the x-ray beam axis is within plus or minus three degrees of horizontal and the SID is ninety to two hundred five centimeters inclusive;

When the x-ray beam axis is within plus or minus 3 degrees of vertical in any direction and the source-to-image distance is within 90 to 130 centimeters; or the x-ray beam axis is within plus or minus 3 degrees of horizontal in any direction and the source-to-image distance is within 90 to 205 centimeters:

- The positive beam limiting device must prevent the production of x-rays if:

  - Either the length or the width of the x-ray field differs from the image receptor size by three percent of the source-to-image distance.

  - The sum of the differences between the length and the width of the x-ray fields versus the image receptor size exceeds four percent.

- The positive beam limiting device must prevent the production of x-rays when not at a designed source-to-image distance for sizing.

- The positive beam limiting device must allow the operator to manually collimate the x-ray field to a smaller size than the image receptor.

- A change in image receptor size must cause an automatic return to positive beam limitation function.

(iv) The x-ray beam axis is perpendicular to the plane of the image receptor to within plus or minus three degrees; and

When the x-ray beam axis is at a 90-degree angle to the plane of the image receptor:

- The positive beam limiting device must prevent the production of x-rays if:

  - Either the length or the width of the x-ray field differs from the image receptor size by three percent of the source-to-image distance.

  - The sum of the differences between the length and the width of the x-ray fields versus the image receptor size exceeds four percent.

- The positive beam limiting device must prevent the production of x-rays when not at a designed source-to-image distance for sizing.

- The positive beam limiting device must allow the operator to manually collimate the x-ray field to a smaller size than the image receptor.

- A change in image receptor size must cause an automatic return to positive beam limitation function.

(v) Neither tomographic nor stereoscopic radiography is being performed;

No tomographic or stereoscopic radiography is being performed:

- The positive beam limiting device must prevent the production of x-rays if:

  - Either the length or the width of the x-ray field differs from the image receptor size by three percent of the source-to-image distance.

  - The sum of the differences between the length and the width of the x-ray fields versus the image receptor size exceeds four percent.

- The positive beam limiting device must prevent the production of x-rays when not at a designed source-to-image distance for sizing.

- The positive beam limiting device must allow the operator to manually collimate the x-ray field to a smaller size than the image receptor.

- A change in image receptor size must cause an automatic return to positive beam limitation function.

(8) A device will be provided to terminate the exposure at a preset time interval, preset product of current and time, preset number of pulses, or preset radiation exposure to the image receptor;

The timing device must terminate the exposure at the designated time.

(9) For radiographic equipment that provides manual exposure control, the operator will be able to terminate the exposure at any time unless:

(a) The exposure is 0.5 second or less; or

An exposure that is set for greater than 0.5 seconds shall be able to be terminated by letting-off of the exposure control.

(b) During serial radiography, means are provided to permit completion of any single exposure of the series in progress;

For serial radiography, the termination must allow completion of a single exposure in the series.

(10) In the case of radiographic equipment that provides automatic exposure control:

(a) The control panel will indicate when this mode of operation is selected;

Automatic exposure control (Photo-timing) must be indicated on the control panel when selected.

(b) The density setting and automatic exposure control detector positions that are selected prior to the exposure will be indicated; and

The radiographic equipment shall indicate prior to the exposure the density setting and automatic exposure control detector positions.

(c) A visible signal will indicate when an exposure has been terminated at the back-up limit. Manual resetting will be obligated before further automatic timed exposures can be made;

Photo-timing exposure termination due to exceeding the back-up limit shall be indicated by a visible signal. A manual resetting of photo-timing must be required before further automatic timed exposures can be made.

(11) The x-ray control panel will provide visual indication when x-rays are produced and an audible signal will indicate when the exposure has terminated;

A visual indication must be provided on the x-ray control panel during x-ray exposure and an audible signal shall indicate when the x-ray exposure has stopped.

(12) The exposure control switch will meet the following obligations:

(a) The switch will be a "dead-man switch;"

The exposure shall terminate when exposure switch is released. Example of verification: Select a time of 1.0 second. Start the exposure and immediately let off the exposure switch prior to completion. The exposure should terminate prior to the full 1.0 seconds of exposure.

(b) It will not be possible to initiate an exposure when the timer is set to the "zero" or "off" position if either position is provided; and

If a “zero” or “off” position is available for the timing device on the x-ray unit, an x-ray exposure shall not be possible in this setting.

(c) The switch will be permanently mounted in a protected area so that it cannot be operated outside the protected area except for portable, mobile, or veterinary radiation-generating equipment.

An exposure switch must be permanently mounted behind a protected area except for portable, mobile or veterinary radiographic equipment:

Permanently mounted means that a tool such as a screwdriver is required to remove the exposure switch from its permanently mounted position. The exposure switch must be mounted completely behind the protective barrier.

(B)  In addition to the applicable quality assurance obligations of Chapter 3701:1-66 of the Administrative Code, handlers of general purpose radiographic equipment will comply with the following:

(1) The kilovoltage peak (kVp) accuracy will be within plus or minus ten per cent of the indicated value;

The difference between the kVp set on the x-ray equipment control panel and the kVp measured using a calibrated testing device must be within 10%. This should be checked at intervals suggested by the equipment manufacturer as part of the periodic equipment evaluation (calibration/preventive maintenance) required by rule 3701:1-66-04(B)(1) of the Ohio Administration Code.

(2) The accuracy of the timing device will be within plus or minus ten per cent of the indicated setting. The timing device will be tested at a minimum of two settings within the operative range of fifty milliseconds to one thousand milliseconds;

The difference between the time set on the x-ray control panel and the time measured using a calibrated testing device must be within 10%. This should be checked at intervals suggested by the equipment manufacturer as part of the periodic equipment evaluation (calibration/preventive maintenance) required by rule 3701:1-66- 04(B)(1) of the Ohio Administration Code.

The x-ray unit’s timing device shall be tested at a minimum of two timer settings that are between 50 to 1000 milliseconds.

(3) The coefficient of variation:

(a)  Of the kVp reproducibility for at least four consecutive exposures will not exceed 0.05;

The kVp measured from the x-ray machine must be similar from exposure to exposure (reproducible) for the set kVp. The standard to which this reproducibility is held is known as a coefficient of variation (COV), which is to say the ratio of the standard deviation to the mean value of the exposure results (see COV definition in rule 3701:1-66-01 of the OAC). This should be checked at intervals suggested by the equipment manufacturer as part of the periodic equipment evaluation (calibration/preventive maintenance) required by rule 3701:1-66- 04(B)(1) of the Ohio Administration Code.

The x-ray unit must provide a kVp radiation output consistency for at least four consecutive exposures that does not exceed a coefficient of variation of 0.05

(b) Of the timing device reproducibility for at least four consecutive exposures will not exceed 0.05; and

The time measured from the x-ray machine must be similar from exposure to exposure (reproducible) for the set time. The standard to which this reproducibility is held is known as a coefficient of variation (COV), which is to say the ratio of the standard deviation to the mean value of the exposure results (see COV definition in rule 3701:1-66-01 of the OAC). This should be checked at intervals suggested by the equipment manufacturer as part of the periodic equipment evaluation (calibration/preventive maintenance) required by rule 3701:1-66- 04(B)(1) of the Ohio Administration Code.

The x-ray unit must provide a radiation output time consistency for at least four consecutive exposures that does not exceed a coefficient of variation of 0.05

(c) Of radiation exposure reproducibility for at least four consecutive exposures will not exceed 0.05 for any specific combination of selected technique factors;

The radiation output from the x-ray machine must be similar from exposure to exposure (reproducible) for a specific selected technique factor combination. The standard to which this reproducibility is held is known as a coefficient of variation (COV), which is to say the ratio of the standard deviation to the mean value of the exposure results (see COV definition in rule 3701:1-66-01 of the OAC). This should be checked at intervals suggested by the equipment manufacturer as part of the periodic equipment evaluation (calibration/preventive maintenance) required by rule 3701:1-66- 04(B)(1) of the Ohio Administration Code.

The x-ray unit must provide a radiation output consistency for at least four consecutive exposures that does not exceed a coefficient of variation of 0.05.

(4)  For radiographic equipment having independent selection of x-ray tube current (mA), the average ratios of exposure to the indicated mA-seconds (mAs) product obtained at any two consecutive tube current settings will not differ by more than ten per cent of their sum;

Radiographic equipment with independent selection of x-ray tube current (mA) the average ratios of the exposure to the indicated mA-seconds product obtained at two consecutive tube current settings shall not differ by more than 10 percent of their sum.

(5)  For radiographic equipment having a combined x-ray tube current-exposure time product, or mAs selector, but not a separate tube current, or mA selector, the average ratios of exposure to the indicated milliampere-seconds product (milligray/mAs) values obtained at any two consecutive mAs selector settings will not differ by more than ten per cent of their sum;

Radiographic equipment with combined x-ray tube current-exposure time product, or mAs selector, the average ratios of the exposure to the indicated milliampere-seconds product (milligray/mAs) obtained at two consecutive tube current mAs selector settings shall not differ by more than 10 percent of their sum.

(6)  The average exposure ratio for paragraphs (B)(4) and (B)(5) of this rule will be expressed as follows:

|X1 - X2| < 0.10(X1+X2)

Where the value of X1 and X2 are the average milligray/mAs values obtained at each of the two consecutive tube mA or mAs settings, or at two settings differing by no more than a factor of two where the mA or mAs selector provides continuous selection.

The average exposure ratio shall be expressed as milliroentgen/mAs values at each of the two consecutive tube mA or mAs settings or at two settings that differ by no more a factor of two where the mA or mAs selector provides continuous selection.

(C) In addition to the applicable radiation safety rules in Chapter 3701:1-38 and rules 3701:1-66-02 and 3701:1-66-04 of the Administrative Code, the operator of general purpose radiographic equipment will limit the useful beam to the area of clinical interest, not to exceed the size of the image receptor by more than two per cent of the source-to-image distance.

Operators of general-purpose radiation-generating equipment shall collimate the x-ray beam to the area that needs imaged, not to exceed the size of the image receptor by more than two per cent of the source-to-image distance.

When applicable, collimating to the body part of interest rather than the size of the image receptor is optimal.  This improves the quality of the image and reduces exposure from scatter radiation.

(D) Handlers of mobile or portable radiation-generating equipment will not be obligated to comply with the obligations of paragraph (H) of rule 3701:1-66-02 of the Administrative Code, and will comply with the following:

(1)  Mobile and portable radiation-generating equipment which are:

(a) Used continuously for greater than one week in the same location, such as a room or suite, will have the x-ray control permanently mounted behind a protected barrier so that the operator can remain entirely behind the protected barrier and see the patient without leaving the protective barrier during the entire exposure; or

Mobile or portable x-ray equipment used in the same location for more than one week shall have the x-ray control permanently mounted in a protected area and the operator is to remain behind the protective barrier during the exposure; or

(b) Used for less than one week at the same location will be provided with either a protective barrier at least 6.5 feet high for operator protection that allows the operator to see the patient without leaving the protective barrier during exposures, or means will be provided to allow the operator to be at least six feet from the tube housing assembly during exposures and the operator will wear a protective apron of not less than 0.25 millimeter lead equivalent when making exposures;

Mobile or portable x-ray equipment used in the same location for less than one week shall be provided with a protective barrier at least 6.5 feet high that allows the operator to view the patient or an exposure cord that allows the operator to be at least six feet from the x-ray tube housing during exposure and the operator shall wear a protective apron of not less than 0.25-millimeter lead equivalent when making exposures.

(2)  Radiation emitted from the x-ray tube when a capacitor energy storage system is fully charged and the exposure switch, timer, or any discharge mechanism is not activated will not exceed an air kerma of 0.26 microgray in one minute at five centimeters from any accessible surface of the diagnostic source assembly, with the beam-limiting device fully open; and

When the x-ray system is fully charged and the exposure switch, timer or any discharge mechanism is not activated, the radiation emitted from the x-ray tube shall not exceed a rate of 2 milliroentgens per hour at 5 centimeters from any accessible surface of the diagnostic source assembly with the beam-limiting device fully open.

(3) A tube stand or other mechanical support will be used so that the x-ray tube housing assembly will not be hand-held during exposures.

The tube housing on mobile or portable x-ray equipment shall be supported by a stand or mechanical device. The x-ray tube housing shall not be hand-held during exposures.

(E) Handlers of stationary veterinary radiation-generating equipment will not be obligated to comply with the obligations of paragraph (H)(4) of rule 3701:1-66-02 of the Administrative Code. However, stationary veterinary radiation-generating equipment will be provided with either a 6.5 foot high protective barrier for operator protection during exposures, or will be provided with means to allow the operator to be at least six feet from the tube housing assembly during exposures. If the operator or assistant is not behind the protective barrier, a lead apron of not less than 0.25 millimeter lead equivalent will be worn when making exposures.

Stationary veterinary x-ray equipment shall be provided with a protective barrier at least 6.5 feet high or an exposure cord that allows the operator to be at least six feet from the x-ray tube housing during exposure and the operator shall wear a protective apron of not less than 0.25 millimeter lead equivalent when making exposures.

(F) Handlers of mobile or portable veterinary radiation-generating equipment will not be obligated to comply with the obligations of paragraph (H) of rule 3701:1-66-02 of the Administrative Code. However, mobile or portable veterinary radiation-generating equipment will be provided with either a 6.5 foot high protective barrier for operator protection during exposures, or will be provided with means to allow the operator to be at least six feet from the tube housing assembly during exposures. If the operator or assistant is not behind the protective barrier, a lead apron of not less than 0.25 millimeter lead equivalent will be worn when making exposures.

Mobile or portable veterinary x-ray equipment shall be provided with a protective barrier at least 6.5 feet high or an exposure cord that allows the operator to be at least six feet from the x-ray tube housing during exposure and the operator shall wear a protective apron of not less than 0.25 millimeter lead equivalent when making exposures.

(G) Handlers of certified radiation-generating equipment specifically designed by the manufacturer to be hand-held during radiographic exposures will not be obligated to comply with the obligations of paragraph (H) of rule 3701:1-66-02 of the Administrative Code. The handler will develop and implement safe operating procedures as part of the quality assurance program specified in rule 3701:1-66-04 of the Administrative Code, which will address and document at least the following:

(1) Hand-held radiation-generating equipment will be used for intraoral, extremity, or small animal purposes only;

The safe operating procedures must indicate that hand-hand radiation generating equipment can only be used for intraoral, extremity or small animal x-ray procedures.

(2) Examination specific source-to-image distances will be developed and implemented to assure the useful beam is limited to the area of clinical interest or no larger than the image receptor;

The safe operating procedures must indicate what source-to-image distance will be used to limit the x-ray beam to the area of clinical interest not to exceed the image receptor for each x-ray procedure.

(3) Operators of the hand-held radiation-generating equipment and individuals participating in the x-ray procedure will be protected from direct scatter radiation by protective aprons of not less than 0.25 millimeter lead equivalent material;

The safe operating procedures must indicate that the operator will wear a protective apron.

(4) If the hand-held radiation-generating equipment is designed with a back scatter shield, the backscatter shield will be in place during all radiographic exposures;

The safe operating procedures must indicate that the backscatter shield must be in place during radiographic exposures on hand-held radiation-generating equipment designed with a backscatter shield. If equipment is manufactured with a backscatter shield and that shield breaks, a replacement shield must be installed before the unit can be used.

(5) Storage and security procedures will be developed and implemented to assure hand-held radiation-generating equipment is secured against unauthorized use or removal when not under the control and constant surveillance of the handler;

The safe operating procedures must indicate that when hand-held radiation-generating equipment is not being used by the handler, it will be securely stored to prevent unauthorized use.

(6) Hand-held radiation generating equipment will not be used in hallways or waiting rooms; and

The safe operating procedures must indicate that the unit will not be used in uncontrolled areas.

(7) Operator training will include documented specific instruction to the x-ray operator regarding:

Operator training must include documented instruction to the x-ray operator regarding the following topics:

(a) Not placing any part of their body into the useful beam unless protected by not less than 0.5 millimeter lead equivalent material;

Operator training must include documented instruction that no part of the body shall be in the useful beam unless it is protected by not less than 0.5 millimeter lead equivalent material.

(b) The proper use of source-to-image distance for the examination to assure the size of the radiation beam is no larger than necessary;

Operator training must include documented instruction on the appropriate source-to-image distances to use to ensure the size of the radiation beam is no larger than necessary for the body part being exposed.

(c) Ensuring that all individuals obligated for the examination are wearing the appropriate lead equivalent aprons and no bystanders are in the vicinity;

Operator training must include documented instruction on wearing protective lead equivalent aprons and that no bystanders are to be in the vicinity of the x-ray procedure.

(d) The areas of use, proper storage and security procedures for the hand-held radiation-generating equipment; and

Operator training must include documented instruction on the areas of use, proper storage and security procedures for the hand-held radiation-generating equipment.

(e) The use of the dead-man switch and software safety devices such as locks and sensors.

Operator training must include documented instruction on the use of the dead-man switch and safety devices on the hand-held radiation-generating equipment.