3701:1-66-08     Mammography radiation-generating equipment.

This rule provides standards for radiation-generating equipment used for screening and diagnostic mammography, and mammography equipment used for invasive localization and stereotactically-guided breast biopsy purposes, except as provided by paragraphs (E) and (F) of this rule. In addition, a handler of mammography radiation-generating equipment that uses either stationary or mobile installations, will comply with all applicable standards in 21 C.F.R. part 1020 (as effective on the effective date of this rule).

These regulations address mammography radiation-generating equipment used for the following:

- Screening;

- Diagnostic;

- Invasive localization;

- Stereotactically-guided breast biopsy; and

- Biopsy of tissue.

(A) In addition to meeting the applicable equipment standards in rule 3701:1-66-02 of the Administrative Code, a handler performing screening or diagnostic mammography will have a valid certificate issued by the U.S. department of health and human services, pursuant to the Mammography Quality Standards Reauthorization Act of 1998, Public Law 105-248, and 21 C.F.R. Part 900 (as effective on the effective date of this rule).

Facilities that perform screening and diagnostic mammography with radiation-generating equipment must have a valid MQSA certificate.  Each mobile unit must have a certificate. The certificate may be a duplicate, if more than one mobile unit is accredited under a single MQSA certificate.

(B) A handler of all types of mammography radiation-generating equipment will comply with the shielding obligations in paragraphs (H)(2) to (H)(4) of rule 3701:1-66-02 of the Administrative Code.

A shielding design and survey by a radiation expert is NOT required, however, structural shielding and protective barriers must be used such that no person other than the patient being examined receives a total effective dose equivalent in excess of the allowable limits in Ohio Administrative Code 3701:1-38-12 and 3701:1-38-13.

The operator shall stand behind, and be able to view the patient, from behind a protective barrier.

(C) In addition to applicable radiation safety obligations in rules adopted pursuant to Chapter 3748. of the Revised Code and rule 3701:1-66-02 of the Administrative Code, a handler of all types of screening and diagnostic mammography radiation-generating equipment will comply with the following:

(1) When a film/screen mammography system is used, clinical films will be processed as soon as possible, but not to exceed twenty-four hours from the time the first clinical image is taken. Facilities utilizing batch processing will:

Facilities that use batch processing must process films within 24 hours.

(a) Use a container to transport clinical films that will protect the film from exposure to light, excessive heat and radiation; and

Facilities that use batch processing must use containers that protect the film from exposure to light, excessive heat and radiation.

(b) Maintain a log to include date and identification of each patient, time of first exposure of each batch, and date and time of each batch processing;

Facilities that use batch processing must maintain a log that includes:

- The examination date and patient identification;

- The time of the first exposure for each batch; and

- The date and time each batch was processed.

While batch processing mainly applies to mobile mammography where transportation and batch processing are common, the requirement applies to all instances where batch processing is performed.

(2) Individuals who perform mammography procedures on human beings will hold an Ohio radiographer license in accordance with Chapter 3701-72 of the Administrative Code and will meet at least one of the following qualifications:

Any individual that operates mammography equipment must be a radiographer licensed in Ohio pursuant to Chapter 4773. of the Revised Code.

(a) Documented evidence of having completed the forty contact hours of training obligated by 21 C.F.R. 900.12(a)(2)(ii); or

Any technologist who does not hold R.T.(M) certification through the American Registry of Radiologic Technologists (ARRT) must complete forty contact hours of training in mammography.

(b) Hold advanced certification in mammography issued by the "American Registry of Radiologic Technologists."

Any technologist who holds an Ohio license and ARRT(M) certification meets the requirement of paragraph (C)(2) of this rule, without the need of further educational courses in mammography.

(D) In addition to all applicable quality assurance obligations in rules 3701:1-66-02 and 3701:1-66-04 of the Administrative Code, the facility will maintain phantom and quality control images for three months.

Records of quality control images shall be maintained for three months.

(E) Radiation-generating equipment designed for mammography, but used exclusively for radiography of tissue from a biopsy, will be exempt from paragraphs (A) to (D) of this rule, and will comply with paragraphs (E), (H)(2), (H)(3) and (J) of rule 3701:1-66-02 of the Administrative Code.

This rule applies ONLY to equipment that is used for ex vivo radiography of excised breast tissue (no live patient exposure).  Typically, these units are de-certified mammography units that are no longer used for screening or diagnostic purposes.  This rule does NOT apply to equipment used for invasive procedures such as invasive localization or stereotactically-guided breast biopsy.

(F) Radiation-generating equipment used for radiography of tissue from a biopsy and equipped with an x-ray tube enclosure designed to exclude personnel from its interior during x-ray generation will be exempt from paragraphs (A) to (E) of this rule, and will comply with the obligations set forth in rule 3701:1-68-06 of the Administrative Code.

Radiation-generating equipment used for tissue biopsy and designed to exclude personnel from its interior (cabinet systems) during exposure are exempt from the requirements of this rule.  This radiation-generating equipment must comply with Ohio Administrative Code 3701:1-68-06.

(G) Quality control testing by a radiation expert will be conducted on mammography radiation-generating equipment used for invasive localization or having stereotactically-guided breast biopsy capability. Quality control testing for stereotactically-guided breast biopsy equipment will follow the "American College of Radiology (ACR) Practice Parameter for the Performance of Stereotactic-Guided Breast Interventional Procedures" (as revised in 2016). This document is available from the "American College of Radiology, 1891 Preston White Drive, Reston, Virginia 20191, telephone (703) 648-8900."

According to ACR Practice Parameter for the Performance of Stereotactic-Guided Breast Interventional Procedures (as revised in 2016), the medical physicist must conduct the following tests at least annually in accordance with the referenced ACR Stereotactic Breast Biopsy Quality Control Manual.

- Collimation assessment: X-ray beam aligned to image receptor and biopsy window aligned with the image field of view;

- Focal spot performance and system limiting resolution: This test is performed using a line pair test pattern tool;

- kVp accuracy and reproducibility;

- Beam quality assessment (half-value layer measurement);

- Automatic exposure control (AEC) system or manual exposure performance assessment;

- Screen speed uniformity (screen-film) or digital receptor uniformity

- Breast entrance exposure;

- Average glandular dose and exposure reproducibility;

- Image quality evaluation;

- Artifact evaluation; and

- Localization accuracy.

(1) The radiation expert will meet the obligations of the aforementioned ACR guideline; and

The radiation expert who conducts quality control testing must meet the requirements of the (ACR) Practice Parameter for the Performance of Stereotactic-Guided Breast Interventional Procedures (as revised in 2016).

Specifically, section III. B. 1. of the ACR guideline references the "ACR Practice Guidelines for the Performance of Screening and Diagnostic Mammography" for medical physicist qualifications. In addition, the radiation expert should have performed at least 1 hands-on stereotactic breast biopsy unit survey under the guidance of a medical physicist qualified to perform such surveys.

(2) The radiation expert will document and verify that the facility is taking proper corrective actions when results of the quality control tests indicate the need.

If a quality control test exceeds a limit or indicates the need for correction action specified by the ACR guidelines, the radiation expert must document and verify that the facility takes the appropriate corrective action to rectify the matter.