3701:1-66-16     Security screening systems.

(A) Definitions for the purpose of this rule:

(1) "Ambient dose equivalent area product (ADAP)" will be determined using the american national standards institute publication "ANSI/HPS N43.17-2009 Radiation Safety for Personnel Security Screening Systems Using X-Ray or Gamma Radiation."

(2) "Interlock" means a device preventing the start or continued operation of equipment that could cause a radiation hazard.

(3) "Mode of operation" means a selectable set of technique factors or machine settings that is pre-determined by the manufacturer for a specific purpose.

(4) "Scan" means the operation necessary to produce one image (e.g., front view) from one radiation source. One radiation source simultaneously producing multiple images also constitutes one scan. Two sources simultaneously producing two images constitute two scans. In some cases several scans may be necessary for a single screening of the subject.

(5) "Screening" means the sum of radiation exposures or scans necessary to image objects concealed on all sides of the body as intended by the system design under normal conditions:

(a) For backscatter systems a screening typically consists of four scans, one from each side.

(b) For transmission systems a screening typically consists of one scan.

(c) For portal systems a screening consists of a complete pass through the inspection zone.

(6) "Security screening system" means radiation-generating equipment used for the sole purpose of screening an individual who is in custody of a law enforcement agency to identify contraband items that would present a security threat within a secured facility perimeter.

(7) "Technique factors" means the x-ray settings including:

(a) The peak kilovoltage applied to the x-ray tube;

(b) The electric current passing through the x-ray tube; and

(c) The scan time.

(B) Security screening systems will meet the following equipment standards:

(1) Indicators that light only when a scan is in process will be provided and clearly visible to all security screening system operators and anyone approaching the restricted area;

Security screeners shall have lights that illuminate when x-rays are being produced.  The lighted indicators must be visible to all security screening operators and anyone approaching the restricted area.

Multiple lighted indicators may be used to meet the requirements of this rule (i.e. one for the operator and another (or multiple others) for anyone approaching the restricted area.

You may wish to reference ANSI/HPS N43.17-2009 for further details.

(2) Power to the system will be controlled by a key switch;

Security screeners shall have a key-actuated control to turn the power on making it ready to generate radiation. The system shall not be able to generate radiation without the key turning the electric power on.

You may wish to reference ANSI/HPS N43.17-2009 for further details.

(3) A device to terminate x-ray exposure at any time during a scan;

Security screeners shall have a device to terminate exposure at any time. You may wish to reference ANSI/HPS N43.17-2009 for further details.

Example: Emergency stop button.

(4) Access panels to x-ray source and detector will be provided with at least one safety interlock;

Interlock means a device preventing the start or continued operation of equipment that could cause a radiation hazard.

Security screeners shall be designed with at least one safety interlock on the access panel. You may wish to reference ANSI/HPS N43.17-2009 for further details.

(5) Operational safety interlocks will terminate the x-ray exposure in the event of any system problem that could result in abnormal or unintended radiation emission;

Security screeners shall be designed with operational safety interlocks that will terminate the exposure in the event of any system problem that could result in abnormal or unintended radiation emission.

This may be software-based interlock systems that sense abnormalities in equipment function such as unintended stoppage of beam motions, abnormal or unintended x-ray source out-put, computer system malfunctions, and shutter or beam stop mechanism malfunction and will terminate the exposure. You may wish to reference section 7.2.1 l. of ANSI/HPS N43.17-2009 for further details.

(6) Following any premature termination, the security screening system will prevent resumption of x-ray generation until the normal control sequence is reset for a new scan;

Security screeners shall have to be reset in order to resume operation following a premature termination.

Example: An operational safety interlock terminates the exposure due to a system malfunction. The operator must now reset the machine in order to begin scanning again.

Example: A worker pressed the emergency stop button during a scan. The operator must now reset the machine in order to begin scanning again.

You may wish to reference section 7.2.1 n. of ANSI/HPS N43.17-2009 for further details regarding the origin of this requirement.

(7) Equipment designed to control the exposure output using multiple modes of operation will indicate the selected mode prior to each scan;

Security screeners capable of operating in both the limited and general use modes shall indicate which mode is selected prior to performing the scan.

Example: Some units have a three-level system, where Level 1 is the general use mode, Level 2 is the default limited use mode, and Level 3 is the excessive BMI limited use mode.  The system indicates the mode by lighting up a dot next to the level.

(8) Technique factors can not be adjustable and will be preset by the manufacturer for each mode of operation;

Security screeners shall have preprogrammed techniques that cannot be adjusted by the operator.

You may wish to reference section 7.2.2 b. or section 7.2.3 b. of ANSI/HPS N43.17-2009 for further details.

(9) A means will be provided to terminate the exposure at a preset time interval or exposure;

Security screening systems shall have the means to terminate the exposure at a preset time interval.

You may wish to reference section 7.2.2 e. or section 7.2.3 e. of ANSI/HPS N43.17-2009 for further details.

(10) When the x-ray tube is operated at its maximum rated tube current for the maximum kilovoltage, the leakage dose will not be greater than 2.5 microsievert (0.25 millirem) in any one hour at any point thirty centimeters from any external surface; and

Security screeners shall be designed to limit leakage radiation to 30 centimeters from the external surface to no more than 0.25 millirem in one hour.

This rule addresses radiation outside the primary beam. For systems that employ a shutter to control exposure (i.e. for systems that normally keep high voltage applied to the x-ray tube at times other than during a scan), this rule also addresses the region of the primary beam while the shutter is closed. You may wish to reference section 6.3 of ANSI/HPS N43.17-2009 for further details regarding the origin of this requirement.

(11) The primary x-ray beam will be attenuated by at least one millimeter of aluminum- equivalent total filtration.

Security screeners shall attenuate low level x-rays that do not contribute to the image and only increase dose to the individual being scanned. The filtration shall be equivalent to one millimeter of aluminum. You may wish to reference section 7.1 of ANSI/HPS N43.17-2009 for further details.

(C) Handlers of security screening systems will comply with the following administrative and radiation safety obligations:

(1) No individual will be exposed to the useful beam unless authorized by a law enforcement agency for security benefit;

Individuals shall only be scanned for security purposes as authorized by a law enforcement agency.

(2) No individual will be exposed to the useful beam for demonstration or frivolous purpose;

In accordance with rule 3701:1-66-16(A)(6) security screening systems are only to be used on individuals, who are in the custody of a law enforcement agency (i.e. inmates, prisoners, or intake of individuals under arrest) to identify contraband items within the secured facility. Security screening systems are not to be used on visitors, employees, witnesses, etc., and in accordance with paragraphs (D)(3) and (D)(4) of this rule, a policy must be in place to prohibit the screening of pregnant individuals or minors, even if in the individual is in the custody of a law enforcement agency.

Individuals shall not be scanned for any reason other than security purposes.

For example: Security screening systems shall not be used for training or demonstration unless it is for actual security during hands-on training.

(3) The individual responsible for radiation protection will ensure that all operators are trained in the safe operation of the security screening systems;

The individuals responsible for radiation protection shall ensure all operators are trained on the security screening system as required by paragraph (D)(5) of the rule.

(4) Any radiation-generating equipment that does not meet the provisions set forth in this rule will not be used to irradiate individuals unless the director determines that the continued use will not pose a radiation risk and arrangements have been made to promptly correct the deficiency;

(5) The handler will follow the manufacturer’s recommended maintenance schedule;

Handlers of security screening systems shall follow the manufacturer’s recommended maintenance schedule as usually found in the equipment operation manual.

(6) Radiation-generating equipment will bear a warning label on the control panel or by the exposure switch which cautions individuals that radiation is produced when it is energized; and

The warning label for security screening systems must be on the control panel or by the exposure switch (it is not acceptable to have the label located by the main power switch).

The use of universal symbols is acceptable for meeting the requirements of this rule.

(7) All position locking, holding, and centering devices on radiation-generating equipment components will function as designed by the manufacturer.

All components for position locking, holding, and centering on the security screening system shall function as designed by the manufacturer.

(D) Handlers of security screening systems will comply with the obligations of rule 3701:1-66-04 of the Administrative Code, except for paragraphs (B)(8), (B)(15), (B)(17), (B)(18)(e) and (C). In addition, the written quality assurance program will include the following:

(1) Policy banning the frivolous use of security screening systems where no security benefit is to be derived;

The written quality assurance program shall include policies on prohibiting using the equipment for non-security reasons.

(2) Policy obligating individuals undergoing screening to be positioned facing away from the source of radiation when using transmission security screening systems;

The written quality assurance program shall include policies requiring the screened individual to face away from the x-ray tube.

Transmission systems are those in which the x-ray source and the image detector are located on opposite sides of the screened individual.

(3) Policy banning the exposure of pregnant women;

The written quality assurance program shall include policies prohibiting the exposure of pregnant individuals.

(4) Policy banning the exposure of minors;

The written quality assurance program shall include policies prohibiting the exposure of minors.

(5) Policy that operator training will follow the topics listed in the "Personnel Training" section of the american national standards institute publication "ANSI/HPS N43.17-2009 Radiation Safety for Personnel Security Screening Systems Using X-Ray or Gamma Radiation;"

The written quality assurance program must identify that the operator training follows the topics listed in the "Personnel Training" section of the American National Standards Institute Publication ANSI/HPS N43.17-2009 Radiation Safety for Personnel Security Screening Systems Using X-Ray or Gamma Radiation.

The ANSI/HPS N43.17-2009 operator training topics can be listed in the written quality assurance program, or the written quality assurance program must identify that the operator training follows the topics in the “Personnel Training” section of the ANSI/HPS N43.17-2009.

Implementation of the operator training is verified by documentation that the training occurred.

Topics for Scanners Capable of Either General-Use and/or Limited-Use Operation:

● 	Types of radiation

● 	Sources and magnitude of common exposures

● 	Radiation units

● 	Time, distance, and shielding

● 	Concept of ALARA

● 	Biological effects

● 	Radiation risks

● 	Operating and emergency procedures

● 	Other safety hazards (e.g. unauthorized disassembly of the system)

●	Physical security procedures to prevent unauthorized use or access

●   Operator awareness and control of inspection zones

● 	How to use survey equipment / dosimetry, if applicable

● 	Rights of declared pregnant workers

● 	Regulatory requirements

● 	Supervised practical operations

Additional Topics for Scanners Capable of Limited-Use Operation:

● 	Basic risk communication concepts and training on the administrative controls associated with ensuring that dose to subjects is limited and on the importance of these administrative controls.

Also, see paragraph (C)(3) of this rule regarding the IRRPs responsibility to ensure that all operators are trained in the safe operation of the security screening system.

(6) For general-use full-body security screening systems capable of delivering a maximum effective dose equivalent less than or equal to 0.1 microsievert (ten microrem) per scan: policies and records to show that administrative controls are applied to limit the number of screenings received by any individual such that the reference effective dose equivalent will not exceed:

(a) 0.25 microseivert (twenty-five microrem) per screening; and

(b) Two hundred fifty microsievert (twenty-five millirem) over any twelve month period;

(7) For limited-use full-body security screening systems capable of delivering a maximum effective dose equivalent greater than 0.1 microsievert (ten microrem) per scan: policies and records to show that administrative controls are applied to limit the number of screenings received by any individual such that the reference effective dose equivalent will not exceed:

(a) Ten microsievert (one millirem) per screening; and

(b) Two hundred fifty microsievert (twenty-five millirem) over any twelve month period;

(8) For general-use partial-body security imaging systems capable of delivering a maximum effective dose equivalent less than or equal to 0.1 microsievert (ten microrem) per scan: policies and records to show that administrative controls will be applied to limit the number of screenings received by any individual, such that:

(a) The ADAP will not exceed 0.03 microsievert per square meter (three microrem per square meter) per scan; and

(b) The total number of scans received at the facility in a twelve month period will not exceed N, where N = seventy-five microsievert per square meter per ADAP (seven thousand five hundred microrem per square meter per ADAP);

(9) For limited-use partial-body security imaging systems capable of delivering a maximum effective dose equivalent greater than 0.1 microsievert (ten microrem) per scan: policies and records to show that administrative controls will be applied to limit the number of screenings received by any individual, such that:

(a) The ADAP will not exceed three microsievert per square meter (three hundred microrem per square meter) per scan; and

(b) The total number of scans received at the facility in a twelve month period will not exceed N, where N = seventy-five microsievert per square meter per ADAP (seven thousand five hundred microrem per square meter per ADAP).

(E) Facility, design, shielding and restricted area obligations:

(1) A clearly marked restricted area will be established. The dose outside of the restricted area will not exceed twenty microseivert (two millirem) in any one hour;

The facility may refer to this as the ‘inspection zone’, as defined in section 2.0 of ANSI/HPS N43.17-2009. Please note the use of the term ‘dose’ in this paragraph. While it is acceptable to demonstrate compliance with this requirement by applying the method described in rule 3701:1-38-13(E), the facility will likely calculate the ‘ambient dose equivalent, H*(d)’, as defined in section 2.0 of ANSI/HPS N43.17-2009 document and ensure this result does not exceed 20 uSv (2 mrem) in any one hour.

(2) A means will be provided for the operator responsible for initiating the scan to maintain a full visual surveillance of the screening and restricted area; and

The facility may refer to the restricted area as the ‘inspection zone’, as defined in section 2.0 of ANSI/HPS N43.17-2009. (Special Note: this requirement was derived from sections 6.2 and 7.2.1f. of the ANSI/HPS document).

The surveillance system may be direct, mirror view, or real-time video.

(3) Engineering or administrative controls will be provided to ensure that individuals do not reenter the scanning area from the exit while x-rays are being produced.

The handler shall assure that unintended individuals do not enter the security screener while x-rays are being produced.

This can be accomplished by but is not limited to locked doors, key card scanners, barricades, postings, safe operating procedures or a combination of the items listed.

(F) A health physicist, a radiation expert or a qualified individual designated by a radiation expert will use the american national standards institute publication "ANSI/HPS N43.17-2009 Radiation Safety for Personnel Security Screening Systems Using X-ray or Gamma Radiation" to determine reference dose equivalent limits, as specified in paragraphs (D)(6) and (D)(7) of this rule, and ADAP, as specified in paragraphs (D)(8) and (D)(9) of this rule, as follows:

The term ‘health physicist’ is not defined in the OAC. The intent of this paragraph is to allow qualified individuals who may not meet the definition of a ‘radiation expert’, because of the scope of their credentials, to perform the radiation measurements and calculations associated with the safe operation of security screening systems.

A ‘radiation expert’ or an individual who holds certification as a health physicist through the American Board of Health Physics (CHP) is qualified.

"Radiation expert" means an individual who meets the qualifications of:

- Applicable paragraphs of rule 3701:1-66-03 of the Administrative Code;

- Paragraph (D) of rule 3701-83-45 of the Administrative Code, for any facility providing radiation therapy services;

- Paragraph (C)(3) of rule 3701-83-52 of the Administrative Code for CT equipment, or paragraph (F)(3) of rule 3701-83-52 of the Administrative Code for fluoroscopy, at any facility providing CT or fluoroscopy services; or

- 21 C.F.R. 900.12(a)(3) (as effective on the effective date of this rule) for any facility providing mammography services.

(1) Upon installation;

All calculations must be performed prior to human exposure.

(2) Annually; and

‘Annually’ means at intervals not to exceed one year, plus or minus one month (found in rule 3701:1- 38-01(A)(18) of the OAC.)

(3) After any maintenance or change that may affect the reference effective dose or ADAP.

Maintenance or change that would necessitate a re-measurement & re-calculation of the reference effective dose (full-body scanners) or ADAP (partial-body scanners) includes, but is not limited to:

- Any adjustment of the technique factors or machine settings that define a ‘mode of operation’ (this may ONLY be done by the manufacturer);

- Any adjustment that effects the radiation output level(s)

- Replacement of the x-ray tube; or

- Replacement of the x-ray detector, if the detector controls the exposure levels (i.e. automatic exposure control) or if any adjustments or calibrations are made which affect the radiation output level(s).

(G) Screening systems capable of delivering an effective dose equivalent greater than ten microsievert (one millirem) per scan will not be used for non-medical screening of human beings for security purposes.

As screening does not necessitate direction by a licensed practitioner, standard ‘medical’ x-ray equipment (i.e. fluoroscopy or radiography) is not appropriate security screening purposes.